SE(13)014A5

ECC Report <SE24WI40>

Compatibility study between MBANSs operating in the 2300 – 2400 MHz, 2400 – 2483.5 MHz, and 2483.5 – 2500 MHz bands and other systems in the same bands or in adjacent bands

Month YYYY (Arial 9pt bold)

(last updated: Month YYYY) (Arial 9pt)

# Executive summary

This report deals with the analysis of the compatibility between Medical Body Area Network Systems (MBANSs) and other systems operating in the same or adjacent frequencies. It has been developed in response to ETSI system reference document TR 101 557, and in accordance with the conclusions of WG FM on the spectrum designation options for MBANSs and further considerations within SRD/MG and SE24.  . The compatibility of MBANSs is studied in three frequency bands: 2300 - 2400 MHz, 2400 – 2483.5 MHz, and 2483.5 - 2500 MHz.

TBD (style: ECC Paragraph)

(advice: the Executive Summary should provide a short and concise explanation on the purpose of the respective ECC Report and should clearly indicate the covered subjects to which it applies. In addition, it should clearly explain the application of the document.)

TABLE OF CONTENTS

0	Executive summary	2

1	Introduction	8

2	ALLOCATIONS IN THE BANDS 2300 – 2400 MHz, 2400 – 2483.5 MHZ, AND 2483.5 – 2500 MHZ	9

3	characteristics of MBANS	11

3.1	Operation Scenario	12

4	Compatibility Study for the Band 2300 – 2400 MHz	15

4.1	Aeronautical Telemetry	15

4.1.1	Healthcare facility MBANSs	16

4.1.1.1	Interference from MBANSs to ATS receivers	16

4.1.1.2	Interference from ATS transmitters to MBANSs	17

4.1.2	Home MBANSs	18

4.1.2.1	Interference from MBANSs to ATS receivers	19

4.1.2.2	Interference from ATS transmitters to MBANSs	20

4.1.3	Ambulance MBANSs	21

4.1.4	Summary Aeronautical Telemetry	21

4.2	Amateur	22

4.2.1	Healthcare facility MBANSs	23

4.2.1.1	Interference from MBANSs to Amateur	24

4.2.1.2	Interference from Amateur to MBANSs	25

4.2.2	Home MBANSs	26

4.2.2.1	Interference from MBANSs to Amateur	27

4.2.2.2	Interference from Amateur to MBANSs	29

4.2.3	Ambulance MBANSs	30

4.2.3.1	Interference from MBANSs to Amateur	31

4.2.3.2	Interference from Amateur to MBANSs	32

4.2.4	Summary Amateur	33

4.3	Broadband wireless systems (BWS)	34

4.3.1	Healthcare facility MBANSs	37

4.3.1.1	Interference from LTE DL to MBANSs	38

4.3.1.2	Interference from MBANSs to LTE DL	39

4.3.1.3	Interference from LTE UL to MBANSs	40

4.3.1.4	Interference from MBANSs to LTE UL	41

4.3.2	Home MBANSs	42

4.3.2.1	Interference from LTE DL to MBANSs	43

4.3.2.2	Interference from MBANSs to LTE DL	44

4.3.2.3	Interference from LTE UL to MBANSs	47

4.3.2.4	Interference from MBANSs to LTE UL	48

4.3.3	Ambulance MBANSs	49

4.3.3.1	Interference from LTE DL to MBANSs	50

4.3.3.2	Interference from MBANSs to LTE DL	50

4.3.3.3	Interference from LTE UL to MBANSs	51

4.3.3.4	Interference from MBANSs to LTE UL	52

4.3.4	Summary BWS	53

4.4	SAP/SAB (ENG/OB) video links	57

4.4.1	Healthcare Facility MBANSs	58

4.4.1.1	Interference from MBANSs to SAP/SAB video links	59

4.4.1.2	Interference from SAP/SAB video links to MBANSs	60

4.4.2	Home MBANSs	61

4.4.2.1	Interference from MBANSs to SAP/SAB video links	62

4.4.2.2	Interference from SAP/SAB video links to MBANSs	63

4.4.3	Ambulance MBANS	64

Interference from MBANSs to SAP/SAB video links	65

4.4.3.1	65

4.4.3.2	Interference from SAP/SAB video links to MBANSs	66

4.4.4	Summary SAP/SAB	67

4.5	ADJACENT Wideband data transmission systems	68

4.5.1	Healthcare facility MBANSs	68

4.5.2	Home MBANSs	69

4.5.3	Summary ADJACENT wideband data transmission systems	69

4.6	ADJACENT Deep Space Research Service (SRS)	69

4.6.1	Healthcare facility MBANSs	70

4.6.2	Home MBANSs	72

4.6.3	Ambulance MBANSs	74

4.6.4	Summary Deep Space SRS	76

5	Compatibility Study for the Band 2400 – 2483.5 MHz	77

5.1	Wideband data transmission systems	77

5.1.1	Healthcare facility MBANSs	78

5.1.1.1	Interference from MBANSs to wideband data transmission systems	78

5.1.1.2	Interference from wideband data transmission systems to MBANSs	79

5.1.2	Home MBANSs	79

5.1.2.1	Interference from MBANSs to wideband data transmission systems	79

5.1.2.2	Interference from wideband data transmission systems to MBANSs	80

5.1.3	Ambulance MBANSs	81

5.1.4	Summary wideband data transmission systems	81

5.2	ISM	81

5.3	Non-specific SRDs	81

5.4	Railway Applications	81

5.5	RFID	82

5.6	Amateur and Amateur Satellite	82

5.6.1	Healthcare facility MBANSs	83

5.6.2	Home MBANSs	84

5.6.3	Ambulance MBANSs	85

5.6.4	Summary Amateur and Amateur Satellite	86

5.7	Adjacent broadband wireless systems (BWS)	86

6	Compatibility Study for the Band 2483.5 – 2500 MHz	87

6.1	Active Medical Implants	87

6.1.1	Healthcare facility MBANSs	88

6.1.1.1	Interference from MBANSs to LP-AMI-P	88

6.1.1.2	Interference from MBANSs to LP-AMI	90

6.1.1.3	Interference from LP-AMI-P to MBANSs	90

6.1.1.4	Interference from LP-AMI to MBANSs	91

6.1.2	Home MBANSs	92

6.1.2.1	Interference from MBANSs to LP-AMI-P	92

6.1.2.2	Interference from MBANSs to LP-AMI	94

6.1.2.3	Interference from LP-AMI-P to MBANSs	94

6.1.2.4	Interference from LP-AMI to MBANSs	96

6.1.3	Ambulance MBANSs	96

6.1.4	Summary Active Medical Implants	96

6.2	ISM	97

6.3	MSS Systems (Globalstar mobile phones)	97

6.3.1	Healthcare facility MBANSs	97

6.3.1.1	Interference from MBANSs to Globalstar MES receivers	98

6.3.2	Home MBANSs	99

6.3.2.1	Interference from MBANSs to Globalstar MES receivers	100

6.3.3	Ambulance MBANSs	101

6.3.3.1	Interference from MBANSs to Globalstar MES receivers	101

6.3.4	Summary MSS Systems	102

6.4	CGC (Complementary Ground Component of MSS systems)	103

6.4.1	Healthcare facility MBANSs	103

6.4.1.1	Interference from MBANSs to Globalstar MES receivers in CGC mode	104

6.4.1.2	Interference from CGC base station to MBANSs	105

6.4.2	Home MBANSs	106

6.4.2.1	Interference from MBANSs to Globalstar MES receivers in CGC mode	107

6.4.2.2	Interference from CGC base station to MBANSs	107

6.4.3	Ambulance MBANSs	108

6.4.3.1	Interference from CGC base station to MBANSs	109

6.4.3.2	Interference from MBANSs to Globalstar MES receivers in CGC mode	110

6.4.4	Summary CGC	110

6.5	RDSS/RNSS (Galileo)	111

6.5.1	Healthcare facility MBANSs	112

6.5.1.1	Interference from MBANSs to RNSS receivers (Galileo)	112

6.5.2	Home MBANSs	113

6.5.2.1	Interference from MBANSs to RNSS receivers (Galileo)	114

6.5.3	Ambulance MBANSs	115

6.5.3.1	Interference from MBANSs to RNSS receivers (Galileo)	115

6.5.4	Summary RDSS/RNSS	116

6.6	SAP/SAB (ENG/OB) video links	117

6.7	ADJACENT Wideband data transmission systems	117

6.7.1	Healthcare facility MBANSs	118

6.7.2	Home MBANSs	118

6.7.3	Summary ADJACENT wideband data transmission systems	119

6.8	Adjacent broadband wireless systems (bws)	119

6.8.1	Healthcare facility MBANSs	119

6.8.2	Home MBANSs	120

6.8.3	Ambulance MBANSs	121

6.8.4	Summary adjacent broadband wireless systems	122

7	Conclusions	123

ANNEX 1: heading (style: ECC annex - heading1)	125

ANNEX 2: List of referenceS	126

LIST OF ABBREVIATIONS

| Abbreviation | Explanation |
| --- | --- |
| AFA | Adaptive Frequency Agility |
| AMPS | Advanced Mobile Phone System |
| ATS | Aeronautical Telemetry Systems |
| AP | Access Point |
| APC | Active Power Control |
| BS | Base Station |
| BWS | Broadband Wireless Systems |
| CEPT | European Conference of Postal and Telecommunications Administrations |
| CGC | Complementary Ground Component |
| C/I | Carrier-to-Interference ratio |
| DL | Downlink |
| ECA | European Common Allocation Table (ERC Report 25) |
| ECC | Electronic Communications Committee |
| e.i.r.p | effective isotropic radiated power |
| ENG/OB | Electronic News Gathering / Outside Broadcasting |
| FDD | Frequency Division Duplex |
| FSL | Free Space Link |
| GSO | Geo-Stationary Orbit |
| GSM | Global System for Mobile Communications |
| I/N | Interference-to-Noise ratio |
| IMT | International Mobile Telecommunications system |
| ISM | Industrial, Scientific and Medical applications |
| LBT | Listen Before Talk |
| LoS | Line-of-Sight |
| LP-AMI | Low Power - Active Medical Implant |
| LTE | Long Term Evolution |
| MBANS | Medical Body Area Network System |
| MCL | Minimum Coupling Loss |
| MES | Mobile Earth Station, a user terminal within MSS |
| MS | Mobile Station, handheld user terminal used within mobile communication system |
| MSS | Mobile Satellite Service |
| NFD | Net Filter Discrimination |
| NLoS | Non-Line-of-Sight |
| NMT | Nordic Mobile Telephony |
| OFDMA | Orthogonal Frequency Division Multiple Access |
| Pfd | Power flux density |
| PPDR | Public Protection and Disaster Relief |
| RDSS | Radiodetermination Satellite Service |
| RNSS | Radionavigation Satellite Service |
| Rx | Receiver |
| SAP/SAB | Services Ancillary to Programme making / Services Ancillary to Broadcasting |
| SRD | Short Range Device |
| TACS | Total Access Communication System |
| TDD | Time Division Duplex |
| Tx | Transmitter |
| UE | User Equipment |
| UL | Uplink |
| UMTS | Universal Mobile Telecommunications System |
| VoIP | Voice over Internet Protocol (IP) |
| WiMAX | Worldwide interoperability for Microwave Access |
| WLAN | Wireless Local Area Network |
| WRC | World Radiocommunication Conference |

# Introduction

[….

Initially Medical Body Area Network Systems (MBANSs) will be mostly deployed in healthcare facilities, such as hospitals or emergency care facilities. However they will later extend into the patient’s home in order to enable home healthcare. In addition, MBANSs are also expected to be used in ambulances for monitoring patient vital signs during patient transportation

It is intended that deployment and usage of MBANSs will be at the direction of healthcare professionals. This restriction applies to MBANS operation in both healthcare facilities and out of healthcare facilities (e.g. at patient’s home)….etc….]

[explanation on 2300-2360 ‘extension’ to be added (history of WI40). Not before SRD/MG M57.

The scope of the studies presented herewith is limited to the compatibility between MBANSs and incumbent applications operating in the 2300 – 2500 MHz frequency range and in the bands adjacent to this range. The studies consider, more specifically, the following frequency bands: (1) 2300 – 2400 MHz, (2) 2400 – 2483.5  MHz, and (3) 2483.5 – 2500 MHz. A total spectrum portion of 40 MHz in those bands is considered for MBANS operation. Furthermore this report explicitly deals with the selection and possibility of combination of parts of the three frequency bands. This ensures that enough spectrum is available for MBANSs, whereas the specificities of different usage scenarios and national variations in the use of these bands are considered.

# ALLOCATIONS IN THE BANDS 2300 – 2400 MHz, 2400 – 2483.5 MHZ, AND 2483.5 – 2500 MHZ

The following table summarizes the allocations in the 2300 – 2500 MHz frequency range according to ERC Report 25 [ ]:

Table 1: Allocations according to ERC Report 25 (ECA Table)

| Frequency Range | Utilisation | ERC/ECC Documentation | European Standard |
| --- | --- | --- | --- |
| 2300 – 2400 MHz | Aeronautical Telemetry | ERC/REC 62-02 | - |
|  | Amateur | - | EN 301 783 |
|  | Mobile Applications | - | - |
|  | SAP/SAB | ERC/REC 25-10 | EN 302 064 |
| 2400 – 2450 MHz | Amateur | - | EN 301 783 |
|  | Amateur satellite | - | - |
|  | ISM | - | - (any ref to CENELEC std?) |
|  | Non-specific SRDs | ERC/REC 70-03 | EN 300 440 |
|  | Radiodetermination applications | ERC/DEC/(01)08 ERC/REC 70-03 | EN 300 440 |
|  | Railway applications | ERC/REC 70-03 | EN 300 761 |
|  | RFID | ERC/REC 70-03 | EN 300 440 |
|  | Wideband data transmission systems | ERC/DEC/(01)07 ERC/REC 70-03 | EN 300 328 |
| 2450 – 2483.5 MHz | ISM | - | - (same as above) |
|  | Non-specific SRDs | ERC/REC 70-03 | EN 300 440 |
|  | Radiodetermination applications | ERC/DEC/(01)08 ERC/REC 70-03 | EN 300 440 |
|  | Railway applications | ERC/REC 70-03 | EN 300 761 |
|  | RFID | ERC/REC 70-03 | EN 300 440 |
|  | Wideband data transmission systems | ERC/DEC/(01)07 ERC/REC 70-03 | EN 300 328 |
| 2483.5 – 2500 MHz | Active medical implants | ERC/REC 70-03 | - |
|  | IMT satellite component | - | - |
|  | ISM | - | -(same as above) |
|  | Mobile applications | - | - |
|  | Mobile satellite applications | ECC/DEC/(07)04 ECC/DEC/(07)05 ERC/DEC/(09)02 | EN 301 441 EN 301 473 |
|  | SAP/SAB | ERC/REC 25-10 | EN 302 064 |

1.	This table from ERC Report 25 is not exhaustive. Some important applications may be missing, such as PPDR and military.

The following table lists typical incumbent services/applications in the 2300 – 2500 MHz frequency range considered in the compatibility studies:

Table 2: Considered services / applications for co-channel compatibility

| MBANS candidate band | Typical services/applications potentially relevant for MBANSs | Comments |
| --- | --- | --- |
| 2300 – 2400 MHz | Aeronautical Telemetry | Co-channel |
|  | Amateur | Co-channel |
|  | Mobile Service Applications (LTE TDD/FDD, WiMAX TDD) | Co-channel |
|  | SAP/SAB (ENG/OB) video links | Co-channel |
| 2400 – 2483.5 MHz | Wideband data transmission systems - WLAN EN 300 328 (IEEE-802.11) | Co-channel |
|  | ISM | Co-channel |
|  | Railway applications | Co-channel |
|  | Non-specific SRDs | Co-channel |
|  | RFID | Co-channel |
|  | Amateur/ amateur satellite (2400 – 2450 MHz) | Co-channel |
| 2483.5 – 2500 MHz | Active medical implants (LP-AMI) | Co-channel |
|  | ISM | Co-channel |
|  | MSS (Globalstar mobile phones) | Co-channel |
|  | CGC (Complementary Ground Component of MSS systems) | Co-channel |
|  | RDSS (Galileo) | Co-channel |
|  | SAP/SAB (ENG/OB) video links | Co-channel |

1.	Some important applications may be missing such as PPDR and military. Detailed information on such systems was not available at the time of writing this report.

# characteristics of MBANS

Medical Body Area Network System (MBANS) is a low power radio system used for the transmission of non-voice data to and from medical devices for the purposes of monitoring, diagnosing and treating patients by duly authorized healthcare professionals.

A MBANS consists of one or more on-body wireless sensors—to simultaneously collect multiple vital sign parameters—and/or medical actuator devices that can communicate with a monitoring device placed on/around (up to 10 meters from) the human body. Implantable devices are not part of MBANSs.

Monitoring devices, in their role of MBANS hub, display and process vital sign parameters from MBANS devices and may also forward them (e.g. to a central nurse station) by using wired or wireless technologies other than MBANSs. MBANS hubs also control MBANS devices for the purpose of providing monitoring, diagnosis and treatment of patients. It is expected that, as most typical configuration, a MBANS hub will be associated to only one patient; in the same fashion as a patient monitor is typically wired up to a single patient today. Two MBANS examples are depicted below.

Figure 1: Typical MBANS examples

Oxygen

sat.

Sensor data

Commands

ECG

Oxygen

sat.

ECG

Sensor data

Commands

a) MBAN with portable patient monitor

-

b

) MBAN with bedside patient monitor

MBANS

HUB

MBANS

HUB

The typical number of sensor or actuator devices that communicate with an MBANS hub is in the range of 1 to 5 in healthcare facilities and ambulances. In patient homes, the typical number of sensors is expected to be lower, ranging from 1 to 3.

In most cases the location of MBANS hubs will be rather static and in the immediate proximity of the patient, as depicted in Figure 1 (part b). This applies to MBANS used in healthcare facilities, ambulances, and the patient’s home. Within healthcare facilities, it is however expected that some patients carry a portable MBANS hub on their body.

In the remainder of the document, and without any loss of generality, MBANSs are depicted in simplified manner consisting of an on-body sensor and an off-body hub.

## Operation Scenario

ETSI TR 101 557 discusses the possible environments in which MBANSs are expected to operate and clusters MBANSs into two differentiated types: ‘healthcare facility MBANSs’ and ‘location independent MBANSs’. The latter refers to ambulance and home monitoring scenarios, which in fact present some use case and parameter differences. Whereas the MBANSs expected to be used in healthcare facilities and ambulances are fairly similar to each other, the MBANSs expected to be used in homes have distinctive requirements, which make them de facto a distinct type of MBANSs. For compatibility study purposes, it was considered most appropriate to treat home MBANSs and ambulance MBANSs as different types of systems, abandoning so the use of the generic term ‘location independent’.

In the present study three types of MBANS, or MBANS use cases, are considered, which are schematically depicted in Figure 2, Figure 3, and Figure 4.

MBANS

sensor

e.i.r.p= 0 dBm

3m

MBANS

hub

Figure 2: MBANSs operating in healthcare facility

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

3m

Healthcare facility MBANSs operate exclusively indoor and inside a healthcare facility, in which several MBANSs are simultaneously used on a subset of patients. The restriction to indoor operation within healthcare facilities may be de-facto enabled by requiring MBANS sensors to periodically check that they are within operating distance of their MBANS hub. When the MBANS hub is not reachable, MBANSs data transmission is stopped. That simple mechanism is part of the light licensing concept proposed in the FCC ruling for MBANs operation in the USA [reference].

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

13 dB building penetration loss

Home 1

10m

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Home 2

<10m

Figure 3: MBANSs operating in patient homes

Home MBANSs operate in patient homes, in which an MBANS hub is installed. The patient is mostly inside the home too. However, in a few cases, the patient can be outdoor (e.g. in the backyard) while still in range of the MBANS hub, as depicted in the right part of Figure 3. Worst case interference analyses within this document consider such possibility. The required range and transmit power for home MBANS is higher than for healthcare facility MBANSs.

< 3 m

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB   penetration loss

Figure 4: MBANS operating in ambulance

Ambulance MBANSs are similar to healthcare facility MBANSs in terms of transmit power, and type and amount of data to be transmitted. Their mobility and the lower shielding from the environment are the main differences. It is expected that only one MBANS be used per ambulance.

An overview of the MBANS parameters used for the present study is presented in the table below. They were selected from the System Reference document ETSI TR 101 557 v1.1.1[ ].

Table 3: MBANS parameters

| Parameter | Healthcare Facility MBANSs | Ambulance MBANSs | Home MBANSs |
| --- | --- | --- | --- |
| Max Tx power (dBm) | 0 |  | 13 |
| Operation environment | Indoor only | Inside vehicle only | Mostly indoor. Outdoor possible. |
| Tx-density (Km-2) | 40 | 5 | 10 |
| Coverage radius (m) | 3 |  | 10 |
| Antenna gain (dBi) | 0 (omnidirectional) |  |  |
| Antenna height (m) | 1.5 |  |  |
| Bandwidth (MHz) | 1 – 5 (here 3 MHz is used) |  |  |
| Noise floor (dBm) | -93.2 |  |  |
| Receiver sensitivity (dBm) | -81.9 |  |  |
| Emission mask | See Figure 5 |  |  |
| Interference criterion | C/I = 15dB |  |  |
| Duty cycle per hour | < 10 % |  | < 2 % |

[explain the densities, from SRdoc and more.]

Figure 5: MBANS emission mask (reference BW 3 MHz)

# Compatibility Study for the Band 2300 – 2400 MHz

The compatibility studies presented in this section are structured in the following way:

- Section 4.x:  Introduction of the radio service / application analysed in that subsection, including all relevant technical parameters for the studies

- Section 4.x.1: Compatibility with healthcare facility MBANSs, including MCL calculations and SEAMCAT simulations

- Section 4.x.2: Compatibility with home MBANSs, including MCL calculations and SEAMCAT simulations

- Section 4.x.3: Compatibility with ambulance MBANSs mode, including MCL calculations and SEAMCAT simulations

- Section 4.x.4: Summary and conclusions for the analysed radio service / application

## Aeronautical Telemetry

Aeronautical Telemetry Systems (ATS) are used for aircraft testing. In an ATS, an aircraft-under-test transmits telemetry data to a ground-based receiving station. The usage of ATS is limited to the duration of aircraft testing, and it is therefore not used for other aircraft communications—such as regular flight control. During operation, the high gain parabolic antenna of the receiving station tracks the position of aircraft in the sky. Therefore any interference arriving over terrestrial path may assume to be falling in the side-lobes of the ATS receiver antenna. Although the deployment of these systems in the 2300 – 2400 MHz spectrum is geographically limited in Europe, it was considered necessary to consider ATS in the compatibility studies.

The parameters of digital ATS used in this study were mainly taken from Recommendation ITU-R M.1459. []. ECC Report 110[ ] and ECC Report 6 [ ] were also consulted.

Table 4: Aeronautical telemetry system parameters

|  | Parameter | Value |
| --- | --- | --- |
| Airborne Transmitter | Bandwidth (MHz) | 5 |
|  | Max Tx power (dBm) | 40 |
|  | Antenna gain (dBi) | 0 (omnidirectional) |
|  | Antenna height (m) | Up to 20000. (10000 assumed) |
| Ground Receiver | Receiver sensitivity (dBm) | -92 |
|  | Antenna gain (dBi) | 29 (tracking dish antenna, 2.44 m diameter) |
|  | Antenna diagram | See Figure 6 |
|  | Antenna height (m) | 20 |
|  | I/N objective (dB) | -3 |

Figure 6: Antenna pattern of aeronautical telemetry receiver (spherical)

### Healthcare facility MBANSs

Table 5: MCL calculation in 2360 – 2400 MHz band between healthcare facility MBANSs and ATS

| Victim characteristics | Units | ATS receiver | MBANSs |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 5 | 3 |
| Receiver noise figure | dB | 2 | 10 |
| Receiver antenna height | m | 20 | 1.5 |
| Receiver antenna gain | dBi | 0 (note1) | 0 |
| Operating frequency | MHz | 2380 | 2380 |
| N, receiver thermal noise | dBm | -104.8 | -99.1 |
| I/N objective | dB | -3 | 0 |
| Interferer’s characteristics |  | MBANSs | ATS transmitter |
| e.i.r.p | dBm | 0 | 40 |
| Bandwidth | MHz | 3 | 5 |
| BW correction factor | dB | 0 | -2.22 |
| NFD (adjacent band interf) | dB | 0 | 0 |
| Wall attenuation | dB | 13 | 13 |
| Antenna height | m | 1.5 | 10000 |
| Minimum path loss | dB | 94.8 | 123.8 |
| Interference distance FSL model | km | 0.55 | 15.61 |

Note 1: The gain in side lobes of high gain directional (parabolic) antennas.

#### Interference from MBANSs to ATS receivers

The considered interference scenario is depicted in Figure 7 and has been simulated in SEAMCAT. The simulation settings and results are summarized in Table 6. In order to achieve compatibility, a protection distance of 0.7 km around ATS receivers is required. Given the extremely low number of ATS receivers in the CEPT area, it is considered feasible to impose a restriction on MBANS to operate only within healthcare facilities that lie beyond the protection distance of 0.7 km.

Figure 7: Interference scenario – healthcare facility MBANS into ATS receiver

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Victim ATS receiver

Wanted signal paths

Interference paths

ATS transmitter

Table 6: Interference from healthcare facility MBANS into ATS- Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): ATS link (aircraft to ground station) |  |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 5 MHz |  |
| VLT → VLR path (Note 1) | FSL (user-defined radius, 300 Km) |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm |  |
| ILT density | 40/km2 |  |
| ILT probability of transmission | 0.1 |  |
| ILT → VLR interfering path | Extended Hata, suburban, indoor→outdoor, below roof |  |
| ILT → VLR positioning mode | Closest interferer |  |
| ILT → VLR protection distance | 0 | 0.7 km |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -76.22 (4.36) | -76.21 (4.35) |
| iRSSunwanted, dBm (Std.dev., dB) | -130.74 (18.43) | -143.59 (12.43) |
| Probability of interference (%) (I/N = -3 dB) | 11.9 | 1.1 |

Note 1: ATS ground station antenna pointing at aircraft transmitter.

An interference probability value around 1% was assumed in SEAMCAT to determine the necessary protection distance.

#### Interference from ATS transmitters to MBANSs

The considered interference scenario is depicted in Figure 8 and has been simulated in SEAMCAT. The simulation settings and results are summarized in Table 7. The interference probability from an aircraft-mounted ATS transmitter is very low. The altitude of aircraft-mounted antennas and the non-directionality of both MBANS antennas and ATS transmitter antennas are the main factors contributing to such very low levels of interference to MBANS operation.

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Wanted ATS receiver

Wanted signal paths

Interference paths

ATS transmitter

Figure 8: Interference scenario – ATS transmitter into healthcare facility MBANS

Table 7: Interference from ATS to healthcare facility MBANS - Settings and results

| Simulation input/output parameters | Settings / Results |
| --- | --- |
| Victim Link (VLK): MBANS |  |
| VLK frequency | 2380 MHz |
| VLK bandwidth | 3 MHz |
| VLT → VLR path | FSL (user-defined radius, 3 m) |
| Interfering Link (ILK): ATS link (aircraft to ground station) |  |
| ILK frequency | 2380 MHz |
| ILK bandwidth | 5 MHz |
| ILT Tx power | 40 dBm (reduced by 13 dB wall attenuation) |
| ILT → VLR interfering path | FSL |
| ILT → VLR positioning mode | None |
| ILT → VLR simulation radius | 0.0 Km (ILT directly above VLR) |
| ILT → VLR number of active transmitters | 1 |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -45.2 (4.42) |
| iRSSunwanted, dBm (Std.dev., dB) | -95.18 (1) |
| Probability of interference (%) (C/I = 15 dB) | 0.0 |

### Home MBANSs

Table 8: MCL calculation in 2360 – 2400 MHz band between home MBANSs and ATS

| Victim characteristics | Units | ATS receiver | MBANSs |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 5 | 3 |
| Receiver noise figure | dB | 2 | 10 |
| Receiver antenna height | m | 20 | 1.5 |
| Receiver antenna gain | dBi | 0 (Note 1) | 0 |
| Operating frequency | MHz | 2380 | 2380 |
| N, receiver thermal noise | dBm | -104.8 | -99.1 |
| I/N objective | dB | -3 | 0 |
| Interferer’s characteristics |  | MBANSs | ATS transmitter |
| e.i.r.p | dBm | 13 | 40 |
| Bandwidth | MHz | 3 | 5 |
| BW correction factor | dB | 0 | -2.22 |
| NFD (adjacent band interf) | dB | 0 | 0 |
| Wall attenuation (Note 2) | dB | 0 (outdoor) | 0 (outdoor) |
| Antenna height | m | 1.5 | 10000 |
| Minimum path loss | dB | 120.8 | 136.8 |
| Interference distance FSL model | km | 11.05 | 69.73 |

Note 1: The gain in side lobes of high gain directional (parabolic) antennas.

Note 2: Worst case configuration assumed, in which MBANS transmitter or receiver is outdoors.

#### Interference from MBANSs to ATS receivers

Wanted signal paths

Interference paths

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Home

13 dB building penetration loss

Victim ATS receiver

ATS transmitter

Figure 9: Interference scenario – home MBANS into ATS receiver

Table 9: Interference from home MBANS into ATS - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): ATS link (aircraft to ground station) |  |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 5 MHz |  |
| VLT → VLR path (Note 1) | FSL (user-defined radius, 300 Km) |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 13 dBm |  |
| ILT density | 10/km2 |  |
| ILT probability of transmission | 0.02 |  |
| ILT → VLR interfering path | Extended Hata, suburban, indoor→outdoor, below roof |  |
| ILT → VLR positioning mode | Uniform Density |  |
| ILT → VLR Number of active transmitters | 1 |  |
| ILT → VLR protection distance | 0 | 1.5 km |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -76.19 (4.35) | -76.19 (4.38) |
| iRSSunwanted, dBm (Std.dev., dB) | -130.92 (15.1) | -143.34 (12.44) |
| Probability of interference (%) (I/N = -3 dB) | 7.6 | 1.0 |

Note 1: ATS ground station antenna pointing at aircraft transmitter.

An interference probability value around 1% was assumed in SEAMCAT to determine the necessary protection distance.

Figure 10: Interference scenario – ATS transmitter into home MBANS

Interference from ATS transmitters to MBANSs

Wanted signal paths

Interference paths

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Home

13 dB building penetration loss

Wanted ATS receiver

ATS transmitter

Table 10: Interference from ATS to Home MBANS - Settings and results

| Simulation input/output parameters | Settings / Results |
| --- | --- |
| Victim Link (VLK): MBANS |  |
| VLK frequency | 2380 MHz |
| VLK bandwidth | 3 MHz |
| VLT → VLR path | FSL (user-defined radius, 10 m) |
| Interfering Link (ILK): ATS link (aircraft to ground station) |  |
| ILK frequency | 2380 MHz |
| ILK bandwidth | 5 MHz |
| ILT Tx power | 40 dBm (reduced by 13 dB wall attenuation) |
| ILT → VLR interfering path | FSL |
| ILT → VLR positioning mode | None |
| ILT → VLR Number of active transmitters | 1 |
| ILT → VLR simulation radius | 0.0 Km (ILT directly above VLR) |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -42.65 (4.44) |
| iRSSunwanted, dBm (Std.dev., dB) | -95.19 (1) |
| Probability of interference (%) (C/I = 15 dB) | 0.0 |

### Ambulance MBANSs

Section 4.1.1 shows that coexistence of healthcare facility MBANSs and ATS is possible on the basis of a protection distance. Ambulance MBANSs use the same transmit power level as healthcare facility MBANSs. However their mobile nature does not appear feasible to reach coexistence based on a protection distance. A deeper analysis is therefore considered unnecessary.

### Summary Aeronautical Telemetry

[

Healthcare facility MBANSs:

According to the results presented above, in worst-case situations—modelled with MCL calculations—the interference distance between Aeronautical Telemetry Systems (ATS) and MBANSs is 0.6 km (MBANS as interferer) and 16 km (ATS as interferer). The simulation-based average-case analysis yielded around 12% probability of interference from MBANSs. The result improves significantly with a separation distance of 0.7 km, yielding around 1% interference probability. In the other coexistence direction, the obtained result is 0% probability of interference from ATS.

Co-channel compatibility between healthcare-facility MBANSs and ATS operating in the 2300 — 2400 MHz band is feasible with a reasonable geographical separation distance (e.g. 0.7 km) between the healthcare facility using MBANSs and the ATS ground station. The number of ATS ground stations is limited and their location is known in many countries. The same characteristics, albeit to a lower degree, apply to healthcare facilities.

Home MBANSs and ambulance MBANSs:

In worst-case situations—modelled with MCL calculations—the interference distance between Aeronautical Telemetry Systems (ATS) and home MBANS lies in the range of 11 km (MBANS as interferer) to 69 km (ATS as interferer). The simulation-based average-case analysis yielded 7.6% probability of interference from home MBANSs. The result improves significantly with a separation distance of 1.5 km, yielding around 1% interference probability. In the other coexistence direction, the obtained result is 0% probability of interference from ATS. Ambulance MBANSs were not explicitly analysed.

Co-channel compatibility between home MBANSs or ambulance MBANSs and ATS operating in the 2300 — 2400 MHz band may be possible in practice but cannot be guaranteed due to the required separation distances from the MBANS interferer (1.5 km in average case, up to 11 km in worst-case situation) in combination with the widespread/mobile characteristics of home and ambulance MBANSs.]

## Amateur

The operational characteristics of amateur stations operating in the 2300 – 2400 MHz range vary significantly. However based on the IARU Region-1 VHF Managers Handbook [] and studies for ECC Report-172 / CEPT FM52 they can be categorised as:

- Weak signal reception of Narrowband Terrestrial and EME (Moonbounce) operation, principally in the sub-band 2320 – 2322 MHz, including propagation beacons.

- Some additional narrowband activity in the 2300 – 2305 MHz range

- Data, multimedia, and TV repeaters (point-to-point links and area systems) in other parts of the band

The balance between the above categories may be subject to change as Broadband Wireless Systems are introduced in the 2300 – 2400MHz band.

For studies in the 2360 – 2400 MHz range, it can be assumed that the Amateur Radio stations in such band are used for receiving signals over terrestrial paths (and are unlikely to involve narrowband systems). Recommendation ITU-R M.1732-1 contains parameter ranges for different usage models and frequency bands. The typical parameters of the multimedia and TV repeater operation modes were chosen to represent Amateur Radio in the 2360 – 2400 MHz range. For frequencies near or below 2320 MHz, narrowband activity and propagation beacon deployments were chosen as representative amateur systems to consider (in addition to multimedia and TV, which are present in the entire 2300 – 2400 MHz range).

Table 11 and Table 12 summarize the parameter value choices. Besides ITU-R M.1732-1 [], ECC Report 172 [ ], document SE7(11)054 [ ], and  ECC Report 064[ ] were also consulted for parameter selection.

Table 11: Amateur Radio systems and parameters relevant in 2300 – 2400 MHz range

|  | Parameter | Value (Multimedia) | Value (ATV Repeater) |
| --- | --- | --- | --- |
| Amateur Transmitter | Bandwidth (MHz) | 0.150 | 6 |
|  | Max Tx power (dBm) | 40 | 43 |
|  | Antenna gain (dBi) | 16 (directional) | 10 (omnidirectional) |
|  | Antenna height (m) | 10 | 30 |
| Amateur Receiver | Receiver sensitivity (dBm) | -110 | -94 |
|  | Antenna gain (dBi) | 16 (directional) | 10 (omnidirectional) |
|  | Antenna height (m) | 10 |  |
|  | C/I objective (dB) | 20 |  |

Note: Typical antennas for Repeaters and Beacons are omnidirectional in Azimuth, 10dB gain in Elevation. For directional antennas (e.g for data/multimedia links), 0dBi is assumed in side-lobes.

Table 12: Amateur Radio systems and parameters relevant in 2300 – 2322 MHz range

|  | Parameter | Value (CW-Morse, SSB etc) | Value (Beacons) |
| --- | --- | --- | --- |
| Amateur Transmitter | Bandwidth (MHz) | 0.0027 | 0.002 |
|  | Tx power (dBm) | 50 | 40 |
|  | Antenna gain (dBi) | 20 (directional) | 10 (omnidirectional) |
|  | Antenna height (m) | 10 | 30 |
| Amateur Receiver | Receiver sensitivity (dBm) | -110 | -110 |
|  | Antenna gain (dBi) | 20 (directional) | 20 (directional) |
|  | Antenna height (m) | 10 |  |
|  | C/I objective (dB) | 20 |  |

Note 1: Propagation Beacons typically transmit on a continuous basis (with periodic ‘callsigns’) for long distance reception and monitoring by amateur narrowband stations equipped with low noise (<1dB NF)  receivers and high gain antennas (typically 20dB).

Note 2: Some narrowband amateur stations may be optimised for EME (Moonbounce) activity and have transmit powers in excess of 30dBW, as noted in ECC Report 172. However this study assumes their antennas are aimed above the horizon and highly directional, thus only their receiver side-lobes need be considered, similar to the more typical terrestrial case.

### Healthcare facility MBANSs

Table 13: MCL calculation in 2300 – 2400 MHz band between healthcare facility MBANSs and Amateur Multimedia / ATV Repeater

| Victim characteristics | Units | Multimedia | MBANSs | ATV Receiver | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 0.150 | 3 | 6 | 3 |
| Receiver noise figure | dB | 2 | 10 | 2 | 10 |
| Receiver antenna height | m | 10 | 1.5 | 10 | 1.5 |
| Receiver antenna gain | dBi | 16 | 0 | 10 | 0 |
| Operating frequency | MHz | 2380 | 2380 | 2380 | 2380 |
| N, receiver thermal noise | dBm | -120.1 | -99.1 | -104.0 | -99.1 |
| I/N objective | dB | -10 | 0 | -10 | 0 |
| Interferer’s characteristics |  | MBANSs | Multimedia | MBANSs | ATV Repeater |
| e.i.r.p | dBm | 0 | 56 | 0 | 53 |
| Bandwidth | MHz | 3 | 0.150 | 3 | 6 |
| BW correction factor | dB | -13 | 0 | 0 | -3 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation | dB | 13 | 13 | 13 | 13 |
| Antenna height | m | 1.5 | 10 | 1.5 | 30 |
| Minimum path loss | dB | 120.1 | 142.1 | 111.0 | 136 |
| Interference distance FSL model | km | 10.10 | 127.16 | 3.58 | 63.66 |

Table 14: MCL calculation in 2300 – 2322 MHz band between healthcare facility MBANSs and Amateur

| Victim characteristics | Units | CW-morse and SSB | MBANSs | Beacon Receiver | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 0.0027 | 3 | 0.002 | 3 |
| Receiver noise figure | dB | 1 | 10 | 1 | 10 |
| Receiver antenna height | m | 10 | 1.5 | 10 | 1.5 |
| Receiver antenna gain | dBi | 20 | 0 | 20 | 0 |
| Operating frequency | MHz | 2320 | 2320 | 2320 | 2320 |
| N, receiver thermal noise | dBm | -138.5 | -99.1 | -139.8 | -99.1 |
| I/N objective | dB | -10 | 0 | -10 | 0 |
| Interferer’s characteristics |  | MBANSs | CW-morse and SSB | MBANSs | Beacon Transmitter |
| e.i.r.p | dBm | 0 | 70 | 0 | 50 |
| Bandwidth | MHz | 3 | 0.0027 | 3 | 0.002 |
| BW correction factor | dB | -30.5 | 0 | -31.8 | 0 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation | dB | 13 | 13 | 13 | 13 |
| Antenna height | m | 1.5 | 10 | 1.5 | 30 |
| Minimum path loss | dB | 125.1 | 156.1 | 125.1 | 136.1 |
| Interference distance FSL model | km | 18.43 | 653.81 | 18.43 | 65.38 |

#### Interference from MBANSs to Amateur

The considered interference scenario is depicted in Figure 11 and has been simulated in SEAMCAT. The simulation settings and results are summarized in Table 13. The interference probability from MBANSs into Amateur is low.

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Amateur receiver

Wanted signal paths

Interference paths

Amateur transmitter

Figure 11: Interference scenario – healthcare facility MBANS into Amateur

Table 15: Interference from healthcare facility MBANS to Amateur Radio - Settings and results

| Simulation input/output parameters | Settings / Results (Multimedia) | Settings / Results (ATV Repeater) | Settings / Results (CW-morse and SSB) | Settings / Results (Beacons) |
| --- | --- | --- | --- | --- |
| Victim Link (VLK): Amateur |  |  |  |  |
| VLK frequency | 2380 MHz | 2320 MHz |  |  |
| VLK bandwidth | 150 kHz | 6 MHz | 2.7 kHz | 2 kHz |
| VLT → VLR path | Extended Hata, urban, outdoor→outdoor, above roof | Extended Hata, suburban, outdoor→outdoor, above roof |  |  |
|  | User-defined radius 5 km | User-defined radius 40 Km |  | User-defined radius 100 Km |
| Interfering Link (ILK): MBANS |  |  |  |  |
| ILK frequency | 2380 MHz |  | 2320 MHz |  |
| ILK bandwidth | 3 MHz |  |  |  |
| ILT Tx power | 0 dBm |  |  |  |
| ILT density | 40/km2 |  |  |  |
| ILT probability of transmission | 0.1 |  |  |  |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, above roof |  |  |  |
| ILT → VLR positioning mode | Closest interferer |  |  |  |
| Simulation results |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -67.43 (11.77) | -88.99 (13.19) | -71.55 (12.97) | -108.72 (17.42) |
| iRSSunwanted, dBm (Std.dev., dB) | -140.22 (16.71) | -133.05 (16.65) | -153.41 (16.66) | -154.78 (16.57) |
| Probability of interference (%) (C/I = 20 dB) | 1.4 | 5.2 | 0.73 | 3.2 |

#### Interference from Amateur to MBANSs

The considered interference scenario is depicted in Figure 12 and has been simulated in SEAMCAT. The simulation settings and results are summarized in Table 14. The interference probability from Amateur into MBANSs is very low.

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Amateur receiver

Wanted signal paths

Interference paths

Amateur transmitter

Figure 12: Interference scenario – Amateur into healthcare facility MBANS

Table 16: Interference from Amateur to healthcare facility MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results (Multimedia) | Settings / Results (ATV Repeater) | Settings / Results (CW-morse and SSB) |  | Settings / Results (Beacons) |
| --- | --- | --- | --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |  |  |  |
| VLK frequency | 2380 MHz |  | 2320 MHz |  |  |
| VLK bandwidth | 3 MHz |  |  |  |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |  |  |  |
| Interfering Link (ILK): Amateur |  |  |  |  |  |
| ILK frequency | 2380 MHz |  | 2320 MHz |  |  |
| ILK bandwidth | 150 kHz | 6 MHz | 2.7 kHz |  | 2 kHz |
| ILT Tx power | 40 dBm (10% probability) | 43 dBm |  | 50 dBm | 40 dBm |
| ILT density | 0.2/km2 | 0.001/km2 |  |  |  |
| ILT probability of transmission | 1 |  |  |  |  |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof | Extended Hata, suburban, outdoor→indoor, above roof |  |  |  |
| ILT → VLR positioning mode | Uniform density |  |  |  |  |
| Simulation results |  |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -45.16 (4.51) | -45.08 (4.58) | -45 (4.41) |  | -44.98 (4.37) |
| iRSSunwanted, dBm (Std.dev., dB) | -302.51(72.25) | -125.76 (12.68) | -112.77 (12.8) |  | -122.67 (12.93) |
| Probability of interference (%) (C/I = 15 dB) | 0.1 | 0.01 | 0.13 |  | 0.03 |

### Home MBANSs

Table 17: MCL calculation in 2300 – 2400 MHz band between home MBANSs and Amateur Multimedia / ATV Repeater

| Victim characteristics | Units | Multimedia | MBANSs | ATV Receiver | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 0.150 | 3 | 6 | 3 |
| Receiver noise figure | dB | 2 | 10 | 2 | 10 |
| Receiver antenna height | m | 10 | 1.5 | 10 | 1.5 |
| Receiver antenna gain | dBi | 16 | 0 | 10 | 0 |
| Operating frequency | MHz | 2380 | 2380 | 2380 | 2380 |
| N, receiver thermal noise | dBm | -120.1 | -99.1 | -104.0 | -99.1 |
| I/N objective | dB | -10 | 0 | -10 | 0 |
| Interferer’s characteristics |  | MBANSs | Multimedia | MBANSs | ATV Repeater |
| e.i.r.p | dBm | 13 | 56 | 13 | 53 |
| Bandwidth | MHz | 3 | 0.150 | 3 | 6 |
| BW correction factor | dB | -13 | 0 | 0 | -3 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation (Note 1) | dB | 0 (outdoor) | 0 (outdoor) | 0 (outdoor) | 0 (outdoor) |
| Antenna height | m | 1.5 | 10 | 1.5 | 30 |
| Minimum path loss | dB | 146.1 | 155.1 | 137.0 | 149.0 |
| Interference distance FSL model | km | 201.54 | 586.02 | 71.42 | 284.35 |

Note 1: Worst case configuration assumed, in which MBANS transmitter or receiver is outdoors.

Table 18: MCL calculation in 2300 – 2320 MHz band between home MBANSs and Amateur

| Victim characteristics | Units | CW-morse and SSB | MBANSs | Beacon Receiver | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 0.0027 | 3 | 0.002 | 3 |
| Receiver noise figure | dB | 1 | 10 | 1 | 10 |
| Receiver antenna height | m | 10 | 1.5 | 10 | 1.5 |
| Receiver antenna gain | dBi | 20 | 0 | 20 | 0 |
| Operating frequency | MHz | 2320 | 2320 | 2320 | 2320 |
| N, receiver thermal noise | dBm | -138.5 | -99.1 | -139.8 | -99.1 |
| I/N objective | dB | -10 | 0 | -10 | 0 |
| Interferer’s characteristics |  | MBANSs | CW-morse and SSB | MBANSs | Beacon Transmitter |
| e.i.r.p | dBm | 13 | 70 | 13 | 50 |
| Bandwidth | MHz | 3 | 0.0027 | 3 | 0.002 |
| BW correction factor | dB | -30.5 | 0 | -31.8 | 0 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation | dB | 0 | 0 | 0 | 0 |
| Antenna height | m | 1.5 | 10 | 1.5 | 30 |
| Minimum path loss | dB | 151.1 | 169.1 | 151.1 | 149.1 |
| Interference distance FSL model | km | 367.66 | 2920.46 | 367.66 | 292.05 |

#### Interference from MBANSs to Amateur

MBANS

sensor

MBANS

hub

e.i.r.p= 13  dBm

13 dB building penetration loss

Home

Amateur receiver

Wanted signal paths

Interference paths

Amateur transmitter

Figure 13: Interference scenario – home MBANS into Amateur

Table 19: Interference from home MBANS to Amateur Radio - Settings and results

| Simulation input/output parameters | Settings / Results (Multimedia) | Settings / Results (ATV Repeater) | Settings / Results (CW-morse and SSB) | Settings / Results (Beacons) |
| --- | --- | --- | --- | --- |
| Victim Link (VLK): Amateur |  |  |  |  |
| VLK frequency | 2380 MHz |  | 2320 MHz |  |
| VLK bandwidth | 150 kHz | 6 MHz | 2.7 kHz | 2 kHz |
| VLT → VLR path | Extended Hata, urban, outdoor→outdoor, above roof | Extended Hata, suburban, outdoor→outdoor, above roof |  |  |
|  | User-defined radius 5 km | User-defined radius 40 km |  | User-defined radius 100 km |
| Interfering Link (ILK): MBANS |  |  |  |  |
| ILK frequency | 2380 MHz |  | 2320 MHz |  |
| ILK bandwidth | 3 MHz |  |  |  |
| ILT Tx power | 13 dBm |  |  |  |
| ILT density | 10/km2 |  |  |  |
| ILT probability of transmission | 0.02 |  |  |  |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, above roof |  |  |  |
| ILT → VLR positioning mode | Uniform density |  |  |  |
| Simulation results |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -67.41 (11.81) | -89.08 (13.14) | -71.54 (13.19) | -109.1 (17.57) |
| iRSSunwanted, dBm (Std.dev., dB) | -140.79 (13.53) | -133.97 (13.42) | -154.26 (13.29) | -155.84 (13.24) |
| Probability of interference (%) (C/I = 20 dB) | 0.4 | 2.7 | 0.2 | 1.0 |

#### Interference from Amateur to MBANSs

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Amateur receiver

Wanted signal paths

Interference paths

Amateur transmitter

Figure 14:  Interference scenario – Amateur into home MBANS

Table 20: Interference from Amateur to home MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results (Multimedia) | Settings / Results (ATV Repeater) | Settings / Results (CW-morse and SSB) | Settings / Results (Beacons) |
| --- | --- | --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |  |  |
| VLK frequency | 2380 MHz |  | 2320 MHz |  |
| VLK bandwidth | 3 MHz |  |  |  |
| VLT → VLR path | FSL (user-defined radius, 10 m) |  |  |  |
| Interfering Link (ILK): Amateur |  |  |  |  |
| ILK frequency | 2380 MHz |  | 2320 MHz |  |
| ILK bandwidth | 150 kHz | 6 MHz | 2.7 kHz | 2 kHz |
| ILT Tx power | 40 dBm (10% probability) | 43 dBm | 50 dBm | 40 dBm |
| ILT density | 0.2/km2 | 0.001/km2 |  |  |
| ILT probability of transmission | 1 |  |  |  |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof | Extended Hata, suburban, outdoor→indoor, above roof |  |  |
| ILT → VLR positioning mode | Uniform density |  |  |  |
| Simulation results |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -42.65 (4.42) | -42.62 (4.44) | -42.4 (4.47) | -42.66 (4.35) |
| iRSSunwanted, dBm (Std.dev., dB) | -315.58 (74.54) | -125.75 (12.7) | -112.68 (12.82) | -290.49 (73.05) |
| Probability of interference (%) (C/I = 15 dB) | 0.1 | 0.01 | 0.1 | 1.0 |

### Ambulance MBANSs

Table 21: MCL calculation in 2300 – 2400 MHz band between ambulance MBANSs and Amateur Multimedia / ATV Repeater

| Victim characteristics | Units | Multimedia | MBANSs | ATV Receiver | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 0.150 | 3 | 6 | 3 |
| Receiver noise figure | dB | 2 | 10 | 2 | 10 |
| Receiver antenna height | m | 10 | 1.5 | 10 | 1.5 |
| Receiver antenna gain | dBi | 16 | 0 | 10 | 0 |
| Operating frequency | MHz | 2380 | 2380 | 2380 | 2380 |
| N, receiver thermal noise | dBm | -120.1 | -99.1 | -104.0 | -99.1 |
| I/N objective | dB | -10 | 0 | -10 | 0 |
| Interferer’s characteristics |  | MBANSs | Multimedia | MBANSs | ATV Repeater |
| e.i.r.p | dBm | 0 | 56 | 0 | 53 |
| Bandwidth | MHz | 3 | 0.150 | 3 | 6 |
| BW correction factor | dB | -13 | 0 | 0 | -3 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation | dB | 7 | 7 | 7 | 7 |
| Antenna height | m | 1.5 | 10 | 1.5 | 30 |
| Minimum path loss | dB | 126.1 | 148.1 | 117 | 142.0 |
| Interference distance FSL model | km | 20.15 | 253.72 | 7.14 | 127.01 |

Table 22: MCL calculation in 2300 – 2320 MHz band between ambulance MBANSs and Amateur

| Victim characteristics | Units | CW-morse and SSB | MBANSs | Beacon Receiver | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 0.0027 | 3 | 0.002 | 3 |
| Receiver noise figure | dB | 1 | 10 | 1 | 10 |
| Receiver antenna height | m | 10 | 1.5 | 10 | 1.5 |
| Receiver antenna gain | dBi | 20 | 0 | 20 | 0 |
| Operating frequency | MHz | 2320 | 2320 | 2320 | 2320 |
| N, receiver thermal noise | dBm | -138.5 | -99.1 | -139.8 | -99.1 |
| I/N objective | dB | -10 | 0 | -10 | 0 |
| Interferer’s characteristics |  | MBANSs | CW-morse and SSB | MBANSs | Beacon Transmitter |
| e.i.r.p | dBm | 0 | 70 | 0 | 50 |
| Bandwidth | MHz | 3 | 0.0027 | 3 | 0.002 |
| BW correction factor | dB | -30.5 | 0 | -31.8 | 0 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation | dB | 7 | 7 | 7 | 7 |
| Antenna height | m | 1.5 | 10 | 1.5 | 30 |
| Minimum path loss | dB | 131.1 | 162.1 | 131.1 | 142.1 |
| Interference distance FSL model | km | 36.77 | 1304.52 | 36.77 | 130.45 |

#### Interference from MBANSs to Amateur

MBANS

sensor

MBANS

hub

e.i.r.p= 0  dBm

7 dB penetration loss

Amateur receiver

Wanted signal paths

Interference paths

Amateur transmitter

Figure 15: Interference scenario – Ambulance MBANS into amateur

Table 23: Interference from ambulance MBANS to Amateur Radio - Settings and results

| Simulation input/output parameters | Settings / Results (Multimedia) | Settings / Results (ATV Repeater) | Settings / Results (CW-morse and SSB) | Settings / Results (Beacons) |
| --- | --- | --- | --- | --- |
| Victim Link (VLK): Amateur |  |  |  |  |
| VLK frequency | 2380 MHz |  | 2320 MHz |  |
| VLK bandwidth | 150 kHz | 6 MHz | 2.7 kHz | 2 kHz |
| VLT → VLR path | Extended Hata, urban, outdoor→outdoor, above roof | Extended Hata, suburban, outdoor→outdoor, above roof |  |  |
|  | User-defined radius 5 km | User-defined radius 40 km |  | User-defined radius 100 km |
| Interfering Link (ILK): MBANS |  |  |  |  |
| ILK frequency | 2380 MHz |  | 2320 MHz |  |
| ILK bandwidth | 3 MHz |  |  |  |
| ILT Tx power | 0 dBm |  |  |  |
| ILT density | 5/km2 |  |  |  |
| ILT probability of transmission | 0.1 |  |  |  |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, above roof |  |  |  |
| ILT → VLR positioning mode | Uniform density |  |  |  |
| Simulation results |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -67.32 (12.04) | -89.1.30 (13.12) | -71.34 (13.12) | -108.96 (17.32) |
| iRSSunwanted, dBm (Std.dev., dB) | -138.10 (12.36) | -131.06 (12.18) | -151.53 (12.03) | -152.77 (12.22) |
| Probability of interference (%) (C/I = 20 dB) | 0.7 | 3.1 | 0.3 | 1.5 |

#### Interference from Amateur to MBANSs

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

7 dB penetration loss

Amateur receiver

Wanted signal paths

Interference paths

Amateur transmitter

Figure 16: Interference scenario – Amateur into ambulance MBANS

Table 24: Interference from Amateur radio to ambulance MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results (Multimedia) | Settings / Results (ATV Repeater) | Settings / Results (CW-morse and SSB) | Settings / Results (Beacons) |
| --- | --- | --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |  |  |
| VLK frequency | 2380 MHz |  | 2320 MHz |  |
| VLK bandwidth | 3 MHz |  |  |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |  |  |
| Interfering Link (ILK): Amateur |  |  |  |  |
| ILK frequency | 2380 MHz |  | 2320 MHz |  |
| ILK bandwidth | 150 kHz | 6 MHz | 2.7 kHz | 2 kHz |
| ILT Tx power | 40 dBm (10% probability) | 43 dBm | 50 dBm | 40 dBm |
| ILT density | 0.2/km2 | 0.001/km2 |  |  |
| ILT probability of transmission | 1 |  |  |  |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof | Extended Hata, suburban, outdoor→indoor, above roof |  |  |
| ILT → VLR positioning mode | Uniform density |  |  |  |
| Simulation results |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -45.21 (4.40) | -45.2 (4.36) | -44.96 (4.41) | -44.91 (4.49) |
| iRSSunwanted, dBm (Std.dev., dB) | -316.64 (73.40) | -119.48 (12.85) | -106.63 (12.92) | -116.4 (12.93) |
| Probability of interference (%) (C/I = 15 dB) | 0.1 | 0.1 | 0.3 | 0.1 |

### Summary Amateur

[

Healthcare facility MBANSs:

According to the results presented above, in worst-case situations—modelled with MCL calculations and using an interference criterion of I/N=-10dB for Amateur systems and I/N=0dB for MBANSs—the interference distance between Amateur and MBANSs is in the range of  4 km – 18 km (MBANS as interferer) and 64 km – 654 km (Amateur as interferer). The four different simulation-based average-case analyses yielded 0.7% to 5.2% probability of interference from MBANSs. In the other coexistence direction, the obtained result is <0.2% probability of interference from Amateur systems.

Average case simulation results suggest co-channel coexistence between Amateur systems in the 2300 –2400 MHz band and healthcare facility MBANSs.

Home MBANSs:

In worst-case situations—modelled with MCL calculations—the interference distance between Amateur and MBANSs is in the range of 71 km – 368 km (MBANS as interferer). The simulation-based average-case analyses yielded 0.2% to 2.7% probability of interference from MBANSs. In the other coexistence direction, the obtained result is ≤1% probability of interference from Amateur systems.

Average case simulation results suggest co-channel coexistence between Amateur systems in the 2300 –2400 MHz band and home MBANSs. However in unfavorable conditions (e.g. when an MBANS is located just outside the patient’s home), the required protection distances between Amateur receivers and home MBANSs transmitters exceed 200 km (applicable to entire the 2300 – 2400 MHz range) and 360 km (applicable to the 2300 – 2320 MHz range). The interference potential should not be disregarded and, hence, coexistence cannot be guaranteed.

Ambulance MBANSs:

In worst-case situations—modelled with MCL calculations—the interference distance between Amateur and MBANSs is in the range of 7 km – 37 km (MBANS as interferer). The simulation-based average-case analyses yielded 0.3% to 3.1% probability of interference from MBANSs. In the other coexistence direction, the obtained result is ≤0.3% probability of interference from Amateur systems.

Average case simulation results suggest co-channel coexistence between Amateur systems in the 2300 –2400 MHz band and ambulance MBANSs.

]

## Broadband wireless systems (BWS)

The first international mobile communication system was the analogue NMT system which was introduced in the Nordic countries in 1981, at the same time analogue AMPS was introduced in North America and TACS and J-TACS deployed worldwide. In Europe, the telecommunication administrations in CEPT initiated the GSM project. GSM is the most popular standard for mobile telephony systems in the world. The third generation mobile system UMTS was designed to enable voice and data services in addition to richer mobile multimedia services, including internet access. It started to be deployed in 2001, and by now has 300+ networks deployed, and the number of UMTS connections is estimated to be over 520 million including HSPA (Q1 2010).

The next steps in the 3GPP standardisation of mobile communications systems are referred to as LTE (3GPP technology) and mobile WiMAX (IEEE technology).  The main target for the evolution of 3G mobile communication is to provide the possibility for significantly higher end user data rates compared to what is achievable with for example the first release of the 3G standard. This includes the possibility for higher peak data rate even more possibility for significantly higher data rates over the entire cell area as well as users at the cell edge. Network quality and performance is becoming a clear competitive advantage, as usage and traffic patterns change. With the rapid uptake of smartphones mobile data traffic is expected to grow 15 times by 2017. This enormous amount of data traffic puts requirements on both optimal performance of mobile systems and on emissions from other coexisting technologies.

The investigated frequency allocation is shown in Figure XX. One of the frequency bands defined by IMT for TDD mobile system is 2300-2400 MHz, which has part of the spectrum in common with the proposed MBANS designation. Furthermore, there are allocations adjacent to IMT bands. LTE-TDD is now gaining market traction in all regions as it is commonly considered in the evolution path of any wireless cellular TDD technology (TD-SCDMA, UTRA-TDD and WiMAX™). Globally, many LTE networks were launched in different countries between December 2009 and June 2012.

Figure 21: Spectrum allocation for IMT bands and the proposed MBANS designation

[diagram needs review]

Most of the LTE parameters considered in this report are based on 3GPP TS 36-series documents, describing the LTE standard, as it is the technology adopted for the utilization of IMT-A framework. More specifically the documents TS 36.101, TS 36.104 and TS 36.942 were used, accompanied by the ITU-R M.2039-2 (11/2010) for sharing studies. SEAMCAT is equipped with an OFDMA module specifically designed for LTE. The basis for the SEAMCAT simulations was the similar studies by ECC Report 172, while some specific choices for the OFDMA modules were made according to  liaison statement ECC PT1(10)128 on the LTE parameters. An overview of LTE parameters is presented in Table 1. The definition used for cell radius in SEAMCAT is illustrated in Figure XX and the patterns of the base station antennas, which are taken from the LTE standard, are shown in  Figure XX.

Table 1: LTE TDD parameters

|  | Parameter | Value |
| --- | --- | --- |
| Base Station (BS) | Bandwidth (MHz) | 10 (Note 1) |
|  | Max Tx power (dBm) | 46 |
|  | Antenna gain (dBi) | 17 |
|  | Antenna height (m) | 30 (Note 2) |
|  | Antenna tilt (deg) | -3 |
|  | Feeder loss (dB) | 3 |
|  | ACLR (dB) | Wide area BS: the least stringent of 45 dB and -15 dBm/MHz. Local area BS: the least stringent of 45 dB and -32 dBm/MHz. Home BS: the least stringent of 45 dB and -50 dBm/MHz. |
|  | Spurious emissions (dBm/MHz) | -30 (beyond 10MHz outside operating band) |
|  | Receiver ACS (dB) | 46 |
|  | Noise figure (dB) | 5 |
|  | Interference criterion I/N (dB) | -6 |
| User Equipment (UE) | Bandwidth (MHz) | 10 (Note 1) |
|  | Max Tx power (dBm) | 23 |
|  | Antenna gain (dBi) | 0 (omnidirectional) |
|  | Antenna height (m) | 1.5 |
|  | ACLR (dB) | 30 (1st adjacent channel) |
|  | Spurious emissions (dBm/MHz) | -30 (beyond 10MHz outside operating band) |
|  | Receiver ACS (dB) | 33 |
|  | Noise figure (dB) | 9 |
|  | Interference criterion I/N (dB) | -6 |
| System | Cell Radius R (m) | 250 (in urban scenarios) 500 (in suburban scenarios) |
|  | Cell Layout | 2 tiers, Tri-sector |
|  | Bandwidth of Resource Block (kHz) | 180 |

Note 1: Other bandwidth values are possible according to the LTE standard. 10 MHz was selected as most representative value.

Note 2: Varies between 10 m - 37.5 m above clutter height. 30 m was selected as most representative value.

Figure 1: Cell Radius as defined in SEAMCAT

- (b)

Figure 2: Tri-sector antenna (a) horizontal and (b) vertical pattern

The total emission mask of the system is taken from the standard TS 36.104 and implemented in SEAMCAT is shown in Figure 3. Each user uses 1/3 of the BS bandwidth resources, based on the recommendations by CEPT ECC STG (SEAMCAT Technical Group), thus we have each time 3 active users per BS.

Figure 3: Unwanted emission mask of LTE TDD

In this section the coexistence of LTE systems and MBANSs is evaluated. Results from both MCL calculations and SEAMCAT simulations are presented. Table XX in Section 4.3.4 gives an overview of all studied interference scenarios and the obtained results, followed by the conclusions.

### Healthcare facility MBANSs

The MCL calculations for minimum interference distances between MBANSs and LTE are shown in Table 2 and Table 3 . Depending on the direction of the LTE link (Uplink, UL or Downlink, DL) the MBANSs are interfering and being interfered either by the Base Station, BS or the User Equipment, UE (mobile phones).

Table 2: MCL calculation in 2360 – 2400 MHz band between healthcare facility MBANSs and LTE DL

| Victim characteristics | Units | LTE DL |  | MBANSs |
| --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 10 |  | 3 |
| Receiver noise figure | dB | 9 |  | 10 |
| Receiver antenna height | m | 1.5 |  | 1.5 |
| Receiver antenna gain | dBi | 0 |  | 0 |
| Operating frequency | MHz | 2380 |  | 2380 |
| N, receiver thermal noise | dBm | -94.8 |  | -99.1 |
| I/N objective | dB | -6 |  | 0 |
| Interferer’s characteristics |  | MBANSs |  | LTE DL |
| e.i.r.p | dBm | 0 |  | 60 |
| Bandwidth | MHz | 3 |  | 10 |
| BW correction factor | dB | 0 |  | -5.2 |
| NFD (adjacent band interf) | dB | 0 |  | 0 |
| Wall attenuation | dB | 0 (LTE UE indoor) | 13 (LTE UE outdoor) | 13 |
| Antenna height | m | 1.5 |  | 30 |
| Minimum path loss | dB | 100.8 | 87.8 | 140.8 |
| Interference distance FSL model | km | 1.10 | 0.24 | 109.4 |

Table 3: MCL calculation in 2360 – 2400 MHz band between healthcare facility MBANSs and LTE UL

| Victim characteristics | Units | MBANSs | LTE UL |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 3 | 10 |
| Receiver noise figure | dB | 10 | 5 |
| Receiver antenna height | m | 1.5 | 30 |
| Receiver antenna gain | dBi | 0 | 14 |
| Operating frequency | MHz | 2380 | 2380 |
| N, receiver thermal noise | dBm | -99.1 | -98.8 |
| I/N objective | dB | 0 | -6 |
| Interferer’s characteristics |  | LTE UL | MBANSs |
| e.i.r.p | dBm | 23 | 0 |
| Bandwidth | MHz | 10 | 3 |
| BW correction factor 5 MHz | dB | -5.2 | 0 |
| NFD (adjacent band interf) | dB | 0 | 0 |
| Wall attenuation | dB | 0 | 13 |
| Antenna height | m | 1.5 | 1.5 |
| Minimum path loss | dB | 129.8 | 105.8 |
| Interference distance FSL model | km | 30.49 | 1.74 |
| Interference distance Hata suburban model (when MBANS is indoor) | km | - | 0.106 |

#### Interference from LTE DL to MBANSs

The distance between MBANSs and the reference Base Station (BS) was arbitrarily chosen as half of the LTE cell radius. A different choice has no significant effect on the results.

The considered interference scenario is depicted in Figure 4 and the simulation settings and results are summarized in Table 4.

It should be noted that for this specific simulation scenario the link from the LTE BS to the LTE UE is not implemented in SEAMCAT. Therefore no power control is used for the LTE link, which leads to the highest possible power level from the interfering LTE BS.

MBANS

MBANS

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Wanted signal paths

Interference paths

LTE UE receiver

LTE BS

Figure 4: Interference scenario – LTE DL into healthcare facility MBANS

Table 4: Interference from LTE DL to healthcare facility MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |
| Interfering Link (ILK): LTE DL (BS to UE) |  |  |
| ILK frequency | 2380 MHz |  |
| ILT → VLR interfering path | Extended Hata, outdoor→indoor, above roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Correlated (interfering BS ref. cell) |  |
| Delta X | 0.125 km (R/2) | 0.250 km (R/2) |
| Delta Y | 0 |  |
| System bandwidth | 10 MHz |  |
| BS maximum allowed power | 46 dBm |  |
| Simulation results |  |  |
| Probability of interference (%) (C/I = 15 dB) | 40.8 | 63.8 |

#### Interference from MBANSs to LTE DL

The considered interference scenario is depicted in Figure 5 and the simulation settings and results are summarized in Table 5. Since the OFDMA technology used by LTE does not drop users, the metric used for measuring the interference is the “average bit rate degradation” in the reference (central) cell. The mapping of signal level to bit rate is a function defined by LTE standard and is different for the UL and for the DL. An example is shown in Figure 6.

MBANS

MBANS

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Wanted signal paths

Interference paths

LTE UE

transmitter

LTE BS

Figure 5: Interference scenario – healthcare facility MBANS into LTE DL

Figure 6: An example of OFDMA bitrate mapping

Table 5: Interference from healthcare facility MBANS to LTE DL - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): LTE DL (BS to UE) |  |  |
| VLK frequency | 2380 MHz |  |
| VLT → VLR path | Extended Hata, outdoor→indoor, above roof |  |
|  | urban | suburban |
| Receiver Noise Figure | 9 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| BS maximum allowed power (Wide Area) | 46 dBm |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm |  |
| ILT density | 40/km2 |  |
| ILT probability of transmission | 0.1 |  |
| ILT → VLR interfering path | Extended Hata, indoor→indoor, below roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Uniform density (origin: victim BS ref. cell) |  |
| Simulation results |  |  |
| Average bit rate degradation (%) (Wide Area) | 0.2 | 0.1 |

#### Interference from LTE UL to MBANSs

As in 4.3.1.1, the distance between MBANSs and the reference Base Station (BS) was chosen as half of the LTE cell radius. The considered interference scenario is depicted in Figure 7 and the simulation settings and results are summarized in Table 6.

MBANS

MBANS

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Wanted signal paths

Interference paths

LTE UE

transmitter

LTE BS

Figure 7: Interference scenario – LTE UL into healthcare facility MBANS

Table 6: Interference from LTE UL to healthcare facility MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |
| Interfering Link (ILK): LTE UL (UE to BS) |  |  |
| ILK frequency | 2380 MHz |  |
| ILT → VLR interfering path | Extended Hata, indoor→indoor, below roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Correlated (interfering BS ref. cell) |  |
| Delta X | 0.125 km (R/2) | 0.250 km (R/2) |
| Delta Y | 0 |  |
| Max allowed power of UE | 23 dBm |  |
| Receiver noise figure (BS) | 5 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| Simulation results |  |  |
| Probability of interference (%) (C/I = 15 dB) | 2.2 | 0.4 |

#### Interference from MBANSs to LTE UL

The considered interference scenario is depicted in Figure 8 and the simulation settings and results are summarized in Table 7.

MBANS

MBANS

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Wanted signal paths

Interference paths

LTE UE transmitter

LTE BS

Figure 8: Interference scenario – healthcare facility MBANS into LTE UL

Ser Equipmentpping apping, here

Table 7: Interference from healthcare facility MBANSs to LTE UL - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): LTE UL (UE to BS) |  |  |
| VLK frequency | 2380 MHz |  |
| VLT → VLR path | Extended Hata, indoor→outdoor, above roof |  |
|  | urban | suburban |
| Receiver Noise Figure | 5 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| Max allowed power of UE | 23 dBm |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm |  |
| ILT density | 40/km2 |  |
| ILT probability of transmission | 0.1 |  |
| ILT → VLR interfering path | Extended Hata, indoor→outdoor, above roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Closest interferer (origin: victim BS ref. cell) |  |
| Simulation results |  |  |
| Average bit rate degradation (%) | 1.8 | 4.7 |

It was regarded as unnecessary to consider the placement of the LTE UE outside, since it the impact of MBAN would be equal or lower.

### Home MBANSs

The MCL calculations for worst-case minimum interference distances are presented in Table 8 and Table 9. Since the home MBANSs might be placed outdoor, it can be seen that they are very vulnerable to the BS interference (LTE DL into MBANSs).

Table 8: MCL calculation in 2360 – 2400 MHz band between home MBANSs and LTE DL

| Victim characteristics | Units | LTE DL |  | MBANSs |
| --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 10 |  | 3 |
| Receiver noise figure | dB | 9 |  | 10 |
| Receiver antenna height | m | 1.5 |  | 1.5 |
| Receiver antenna gain | dBi | 0 |  | 0 |
| Operating frequency | MHz | 2380 |  | 2380 |
| N, receiver thermal noise | dBm | -94.8 |  | -99.1 |
| I/N objective | dB | -6 |  | 0 |
| Interferer’s characteristics |  | MBANSs |  | LTE DL |
| e.i.r.p | dBm | 13 |  | 60 |
| Bandwidth | MHz | 3 |  | 10 |
| BW correction factor | dB | 0 |  | -5.2 |
| NFD (adjacent band interf) | dB | 0 |  | 0 |
| Wall attenuation | dB | 0 (Note 1) | 13 | 0 (outdoor) |
| Antenna height | m | 1.5 |  | 30 |
| Minimum path loss | dB | 113.8 | 100.8 | 153.8 |
| Interference distance FSL model | km | 4.93 | 1.1 | 469.7 |

Note 1: Worst case configuration, in which MBANS transmitter and UE are both indoors or outdoors.

Table 9: MCL calculation in 2360 – 2400 MHz band between home MBANSs and LTE UL

| Victim characteristics | Units | MBANSs | LTE UL |  |
| --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 3 | 10 |  |
| Receiver noise figure | dB | 10 | 5 |  |
| Receiver antenna height | m | 1.5 | 30 |  |
| Receiver antenna gain | dBi | 0 | 14 |  |
| Operating frequency | MHz | 2380 | 2380 |  |
| N, receiver thermal noise | dBm | -99.1 | -98.8 |  |
| I/N objective | dB | 0 | -6 |  |
| Interferer’s characteristics |  | LTE UL | MBANSs |  |
| e.i.r.p | dBm | 23 | 13 |  |
| Bandwidth | MHz | 10 | 3 |  |
| BW correction factor | dB | -5.2 | 0 |  |
| NFD (adjacent band interf) | dB | 0 | 0 |  |
| Wall attenuation | dB | 0 (Note 1) | 0 (both outdoor or indoor) | 13 (complementary cases) |
| Antenna height | m | 1,5 | 1,5 |  |
| Minimum path loss | dB | 116.8 | 131.8 | 118.8 |
| Interference distance FSL model | km | 6.97 | 39.17 | 8.73 |

Note 1: Worst case configuration assumed, in which MBANS patient is outdoors.

#### Interference from LTE DL to MBANSs

Scenario is depicted in Figure 9 and simulation settings are in Table 10. Although home MBANSs receive higher wanted signal levels than healthcare facility MBANSs, the interference from BS remains very high.

MBANS

MBANS

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Wanted signal paths

Interference paths

LTE UE receiver

LTE BS

Figure 9: Interference scenario – LTE DL into home MBANS

Table 10: Interference from LTE DL to home MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 10 m) |  |
| Interfering Link (ILK): LTE DL (BS to UE) |  |  |
| ILK frequency | 2380 MHz |  |
| ILT → VLR interfering path | Extended Hata, outdoor→indoor, above roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Correlated (interfering BS ref. cell) |  |
| Delta X | 0.125 km (R/2) | 0.250 km (R/2) |
| Delta Y | 0 |  |
| System bandwidth | 10 MHz |  |
| BS maximum allowed power | 46 dBm |  |
| Simulation results |  |  |
| Probability of interference (%) (C/I = 15 dB) | 27.4 | 50.8 |

#### Interference from MBANSs to LTE DL

The considered interference scenarios are in Figure 10 and Figure 11, and the simulation settings are in Table 11.

MBANS

MBANS

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Wanted signal paths

Interference paths

LTE UE

Receiver 
(indoor)

LTE BS

LTE UE

Receiver 
(outdoor)

Figure 10:  Interference scenario – home MBANS (indoors) into LTE DL

Figure 11: Interference scenario – home MBANS (outdoors) into LTE DL

MBANS

MBANS

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Wanted signal paths

Interference paths

LTE UE receiver (outdoor)

LTE BS

Table 11: Interference from home MBANSs to LTE DL - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): LTE DL (BS to UE) |  |  |
| VLK frequency | 2380 MHz |  |
| VLT → VLR path | Extended Hata, outdoor→indoor/outdoor, above roof |  |
|  | urban | suburban |
| Receiver Noise Figure | 9 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| BS maximum allowed power (Wide Area) | 46 dBm |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 13 dBm |  |
| ILT density | 10/km2 |  |
| ILT probability of transmission | 0.02 |  |
| ILT → VLR interfering path | Extended Hata, indoor→indoor/outdoor, outdoor→ outdoor, below roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Uniform density (origin: victim BS ref. cell) |  |
| Simulation results |  |  |
| Average bit rate degradation (%) | <0.1 (all cases) |  |

For this scenario considered in this subsection it was regarded necessary to also analyse the compatibility situation in the case that MBANSs and LTE are operated in apartments.The scenario when MBANS are applied to a person living in an apartment is quite similar to the case when the person is living in a house. The possible interference scenarios internal to the flat (house) and to and from the outside are considered the same as in the case of a house. The difference, having neighbors above and beneath, was investigated using MCL calculations. The propagation model used for an apartment building was ITU-R P.1238-7.

The MCL calculations indicate that a UE which is 2 levels away need 6.9 meters distance from the MBANS in order to function properly. This would mean a possible interference in a room located 2 levels away straight under/over the MBANS.

For an apartment located adjacent under/over the MCL calculation indicate that a distance of ~ 69 meters is needed. This indicates that there will be interference from MBANS towards UE.

In the other direction, from the UE to MBANS, the interference is lower. However, for an apartment located adjacent under/over the MCL calculation indicate that a distance of ~ 20 meters is needed. This in turn indicates that there will be interference from the UE towards MBANS, at least in the case of co-channel.

|  | Minimum path loss dB | Distance based on ITU-R P.1238-7 (km) N (number of floors) = 1 | Distance based on ITU-R P.1238-7 (km) N (number of floors) = 2 | Distance based on ITU-R P.1238-7 (km) N (number of floors) = 3 |
| --- | --- | --- | --- | --- |
| MBANS interfere Co-ch | 119 | 0.0689 | 0.0069 | 0.0007 |
| UE interfere Co-ch | 101 | 0.0201 | 0,002 | 0.0002 |

#### Interference from LTE UL to MBANSs

Scenario is depicted in Figure 12 and simulation settings are in Table 12.

MBANS

MBANS

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Wanted signal paths

Interference paths

LTE UE

transmitter

LTE BS

Figure 12: Interference scenario – LTE UL into home MBANS

Table 12: Interference from LTE UL to home MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 10 m) |  |
| Interfering Link (ILK): LTE UL (UE to BS) |  |  |
| ILK frequency | 2380 MHz |  |
| ILT → VLR interfering path | Extended Hata, indoor→indoor, below roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Correlated (interfering BS ref. cell) |  |
| Delta X | 0.125 km (R/2) | 0.250 km (R/2) |
| Delta Y | 0 |  |
| Max allowed power of UE | 23 dBm |  |
| Receiver noise figure (BS) | 5 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| Simulation results |  |  |
| Probability of interference (%) (C/I = 15 dB) | 2.2 | 0.1 |

[text referring to explanation about apartments in section 4.3.2.2]

|  | Minimum path loss dB | Distance based on ITU-R P.1238-7 (km) N (number of floors) = 1 | Distance based on ITU-R P.1238-7 (km) N (number of floors) = 2 | Distance based on ITU-R P.1238-7 (km) N (number of floors) = 3 |
| --- | --- | --- | --- | --- |
| UE interfere Co-ch | 101 | 0.0201 | 0,002 | 0.0002 |

#### Interference from MBANSs to LTE UL

For this scenario it was considered necessary to study two cases, with the home MBANS located both indoors and outdoors. Figure 13 and Figure 14 depict these cases. The simulation settings and the results are shown in Table 13.

MBANS

MBANS

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Wanted signal paths

Interference paths

LTE UE transmitter

LTE BS

Figure 13: Interference scenario – home MBANS (indoors) into LTE UL

Figure 14: Interference scenario – home MBANS (outdoors) into LTE UL

MBANS

MBANS

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Wanted signal paths

Interference paths

LTE UE transmitter

LTE BS

Table 13: Interference from home MBANS to LTE UL - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): LTE UL (UE to BS) |  |  |
| VLK frequency | 2380 MHz |  |
| VLT → VLR path | Extended Hata, indoor→outdoor, above roof |  |
|  | urban | suburban |
| Receiver Noise Figure | 5 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| Max allowed power of UE | 23 dBm |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 13 dBm |  |
| ILT density | 10/km2 |  |
| ILT probability of transmission | 0.02 |  |
| ILT → VLR interfering path | Extended Hata, indoor/outdoor →outdoor, above roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Uniform density (origin: victim BS ref. cell) |  |
| Simulation results |  |  |
| Average bit rate degradation (%) (indoor/outdoor) | 1.4 / 6.1 | 6.4 / 19.5 |

### Ambulance MBANSs

Table 14: MCL calculation in 2360 – 2400 MHz band between ambulance MBANSs and LTE

| Victim characteristics | Units | LTE DL | LTE UL | MBANSs |
| --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 10 | 10 | 3 |
| Receiver noise figure | dB | 9 | 5 | 10 |
| Receiver antenna height | m | 1.5 | 30 | 1.5 |
| Receiver antenna gain | dBi | 0 | 14 | 0 |
| Operating frequency | MHz | 2380 | 2380 | 2380 |
| N, receiver thermal noise | dBm | -94.8 | -98.8 | -99.1 |
| I/N objective | dB | -6 | -6 | 0 |
| Interferer’s characteristics |  | MBANSs | MBANSs | LTE UL |
| e.i.r.p | dBm | 0 | 0 | 23 |
| Bandwidth | MHz | 3 | 3 | 10 |
| BW correction factor | dB | 0 | 0 | -5.2 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 |
| Wall attenuation | dB | 7 | 7 | 7 |
| Antenna height | m | 1.5 | 1.5 | 1.5 |
| Minimum path loss | dB | 93.8 | 111.8 | 109.8 |
| Interference distance FSL model | km | 0.49 | 3.92 | 3.11 |
| Interference distance Hata suburban model | km | - | 0.152 | - |
| Interference distance IEEE 802 model C | km | tbd | - | 1.999 |

#### Interference from LTE DL to MBANSs

As shown in section 4.3.1.1, the interference caused by cellular base stations to MBANSs inside healthcare facilities is very high. Taking into account that an ambulance has less attenuation than a hospital, and since all the other parameters remain the same as in the healthcare facility case, the interference probability is expected to be even higher. Thus, it was considered unnecessary to further study this scenario.

#### Interference from MBANSs to LTE DL

The scenario is depicted in Figure 15 and the parameters and results in Table 13. Due to the low density of ambulances with MBANSs inside, the resulting interference is effect is negligible.

Figure 15: Interference scenario – ambulance MBANS into LTE DL

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB   penetration loss

Wanted signal paths

Interference paths

Victim LTE
receiver

LTE BS

Table 15: Interference from ambulance MBANS to LTE DL - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): LTE DL (BS to UE) |  |  |
| VLK frequency | 2380 MHz |  |
| VLT → VLR path | Extended Hata, outdoor→outdoor, above roof |  |
|  | urban | suburban |
| Receiver Noise Figure | 9 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| Max allowed power of UE | 46 dBm |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm |  |
| ILT density | 5/km2 |  |
| ILT probability of transmission | 0.1 |  |
| ILT → VLR interfering path | Extended Hata, indoor →outdoor, below roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Uniform density (origin: victim BS ref. cell) |  |
| Simulation results |  |  |
| Average bit rate degradation (%) (MBANS in ambulance) | 0.0 | 0.1 |

#### Interference from LTE UL to MBANSs

Due to low density of MBANSs and the way SEAMCAT works, an area as small as an ambulance is expected to have no active LTE transmitters inside the ambulance. This is reasonable, since in any case this can be controlled by the personnel of the ambulance. As for the interference coming from LTE UE outside the ambulance, the scenario is shown in Figure 17 and the parameters and results in Table 16.

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB   penetration loss

Wanted signal paths

Interference paths

Interfering LTE UE

LTE BS

Figure 17: Interference scenario – LTE UL into ambulance MBANS

Table 16: Interference from LTE UL to healthcare facility MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |
| Interfering Link (ILK): LTE UL (UE to BS) |  |  |
| ILK frequency | 2380 MHz |  |
| ILT → VLR interfering path | Extended Hata, outdoor→indoor, below roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Correlated (interfering BS ref. cell) |  |
| Delta X | 0.125 km (R/2) | 0.250 km (R/2) |
| Delta Y | 0 |  |
| Max allowed power of UE | 23 dBm |  |
| Receiver noise figure (BS) | 5 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| Simulation results |  |  |
| Probability of interference (%) (C/I = 15 dB) | 8.2 | 1.1 |

#### Interference from MBANSs to LTE UL

The scenario is depicted in Figure 16 and the parameters and results in Table 17.

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB   penetration loss

Wanted signal paths

Interference paths

LTE UE

Victim LTE BS

Figure 16: Interference scenario – ambulance MBANS into LTE UL

Table 17: Interference from ambulance MBANS to LTE UL - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): LTE UL (UE to BS) |  |  |
| VLK frequency | 2380 MHz |  |
| VLT → VLR path | Extended Hata, outdoor→outdoor, above roof |  |
|  | urban | suburban |
| Receiver Noise Figure | 5 dB |  |
| ILK system bandwidth | 10 MHz |  |
| Max subcarriers per BS | 51 |  |
| Number subcarriers per UE | 17 |  |
| Max allowed power of UE | 23 dBm |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm |  |
| ILT density | 5/km2 |  |
| ILT probability of transmission | 0.1 |  |
| ILT → VLR interfering path | Extended Hata, indoor →outdoor, above roof |  |
|  | urban | suburban |
| ILT → VLR positioning mode | Uniform density (origin: victim BS ref. cell) |  |
| Simulation results |  |  |
| Average bit rate degradation (%) (MBANS indoor) | 0.2 | 0.9 |

### Summary BWS

|  | No. | Interference scenario description | Section | MCL distance | Simulation result (urban/suburban) |
| --- | --- | --- | --- | --- | --- |
| Healthcare Facility | 1 | From MBANSs to LTE UL (BS) – MBANSs indoor | 4.3.1.4 | 1.74 km (FSL) 0.106 (Hata, suburban) | 1.8% / 4.7% bitrate degradation 59.3% interference probability (34.5%, if C/I=19dB is used) |
|  | 2 | From MBANSs to LTE DL (UE) – UE and MBANSs indoor | 4.3.1.2 | 1.10 km (FSL) | 0.2% / 0.1% bitrate degradation Ericsson result: 100% interference probability (49.2%, if C/I=19dB is used) |
|  | 3 | From LTE DL (BS) to MBANSs – MBANSs indoor | 4.3.1.1 | 109.4 km (FSL) | 40.8% / 63.8% interference probability |
|  | 4 | From LTE UL (UE) to MBANSs – UE and MBANSs indoor | 4.3.1.3 | 30.49 km (FSL) | 2.2% / 0.4% interference probability |
| Home | 5 | From MBANSs to LTE UL (BS) – UE and MBANSs indoor | 4.3.2.4 | 39.17 km (FSL) | 1.4% / 6.4% bitrate degradation 85% interference probability(48.8%, if C/I=19dB is used) |
|  | 6 | From MBANSs to LTE UL (BS) – UE indoor. MBANSs outdoor | 4.3.2.4 | 8.73 km (FSL) | 6.1% / 19.5% bitrate degradation 99% interference probability(66.3%, if C/I=19dB is used) |
|  | 7 | From MBANSs to LTE DL (UE) – UE and MBANSs indoor | 4.3.2.2 | 4.93 km (FSL) | 0.1% / 0.1% bitrate degradation Ericsson result: 100% interference probability(71.9%, if C/I=19dB is used) |
|  | 8 | From MBANSs to LTE DL (UE) – UE and MBANSs outdoor | 4.3.2.2 | 4.93 km (FSL) | 0.1% 0.8% .1.1n sectionNSs/ 0.1% bitrate degradation Ericsson result: 83% interference probability(15.2%, if C/I=19dB is used) |
|  | 9 | From MBANSs to LTE DL (UE) – UE outdoor and MBANSs indoor | 4.3.2.2 | 1.10 km (FSL) | 0.1% / 0.1% bitrate degradation Ericsson result: 70% interference probability(8.1%, if C/I=19dB is used) |
|  | 10 | From MBANSs to LTE DL (UE) – UE and MBANSs in apartments | 4.3.2.2 | Ericsson result: 0.0069 km (P.1238-7) (probably I/N=0dB used. Should be I/N=-6dB) | n.a. |
|  | 11 | From LTE UL (UE) to MBANSs – UE and MBANSs in apartments | 4.3.2.3 | 0.002 km (P.1238-7) | n.a. |
|  | 12 | From LTE DL (BS) to MBANSs – MBANSs indoor | 4.3.2.1 | 109.4 km (FSL) | 27.4% / 50.8% interference probability |
|  | 13 | From LTE UL (UE) to MBANSs – UE and MBANSs indoor | 4.3.2.3 | 6.97 km (FSL) | 2.2% / 0.1% interference probability |
| Ambulance | 14 | From MBANSs to LTE UL (BS) – MBANSs in ambulance | 4.3.3.4 | 3.92 km (FSL) 0.152 km (Hata) | 0.2% / 0.9% bitrate degradation |
|  | 15 | From MBANSs to LTE DL (UE) – UE outdoor. MBANSs in ambulance | 4.3.3.2 | 0.49 km (FSL) Ericsson: 0.046 km (IEEE 802 model C) | 0.0% / 0.1% bitrate degradation |
|  | 16 | From LTE DL (BS) to MBANSs – MBANSs indoor | 4.3.3.1 | >109.4 km (FSL) 1.047 km (Hata) | >40.8% / >63.8% interference probability |
|  | 17 | From LTE UL (UE) to MBANSs – UE outdoor. MBANSs in ambulance | 4.3.3.3 | 3.11 km (FSL) 0.199 km (Hata) | 8.2% / 1.1% interference probability |

[Explanation of generic simulations in the summary, introduced as worst-case consideration:

If the LTE system is modeled as OFDMA, and UE is the victim, we cannot control where UE will be deployed and/or how far MBANs will be dropped from UE. Simulating the LTE system as generic, we can obtain the probability of a UE being interfered when a MBAN is transmitting close to a UE (i.e. 6-20m)]

In the simulations shown above, the distance between the interferer and the victim may be considerably larger than the reference cell size. This setting represents a larger geographical area, where the interference in the reference cell is the sum from all the interfering devices. However, since the probability of being in close range is small, the results represent an average of the interference situation.

In order to estimate an interference situation using simulations, the distance between the interferer and the victim has to be controlled, such that victim and receiver are in close proximity. This assumption is taken into account in a second set of simulations. The exact parameters for these simulations are given in Appendix X. (Maybe the distances should be stated, for clarity?)

/*START OF KEY PARTS OF ERICSSON CONTRIBUTION (for reference)*/

| Victim |  | LTE BS |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Interferer |  | MBANSs (healthcare) | MBANSs (patient home) | MBANSs (patient flat) | MBANSs (ambulance) |
| Distance based on FSL model (km) | Co-ch | 1.98 | 39.6 | 8.88 | 3.96 |
|  | Adj-ch | 0.199 | 3.96 | 0.88 | 0.39 |
| Distance based on Extended sub-urban Hata model (km) | Co-ch | 0.106 | 0.58 | 0.24 | 1.57 |
|  | Adj-ch | 0.028 | 0.15 | 0.067 | 0.042 |
| ITU-R P.1238-7 | Co-ch | - | - | 0.35 | - |
|  | Adj-ch | - | - | 0.096 | - |

All these results need correction to ACLR of -25dB (instead of -20dB)

| Victim |  | LTE UE |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Interferer |  | MBANSs (healthcare) | MBANSs (patient home) | MBANSs (patient flat) | MBANSs (ambulance) |
| Distance based on FSL model (km) | Co-ch | 0.125 | 2.5 | - | 0.25 |
|  | Adj-ch | 0.013 | 0.25 | - | 0.025 |
| Distance based on Extended sub-urban Hata model (km) | Co-ch | 0.022 | 0.12 | - | 0.032 |
|  | Adj-ch | 0.006 | 0.03 | - | 0.008 |
| Distance based on IEEE802(model C) model (km) | Co-ch | 0.031 | 0.174 | - | 0.046 |
|  | Adj-ch | 0.008 | 0.046 | - | 0.0126 |
| ITU-R P.1238-7 | Co-ch | - | - | 0.0069 | - |
|  | Adj-ch | - | - | 0.0013 | - |

All these results need correction to I/N=-6dB

The ADJACENT results need correction to ACLR of -25dB (instead of -20dB)

| Interferer | Victim | Average Bit rate Loss [%] |  | Probability of Interf. [%] |  | Scenarios |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Co-ch | Adj-ch | Co-ch | Adj-ch |  |
| MBANS healthcare | UE (indoor) | * | * | 100 | 86,7 | N/A |
|  | UE (outdoor) | ** | ** | ** | ** | 3 |
|  | eNB (outdoor) | 4,5 | 0,3 | 59,3 | 40 | 1 |
| MBANS home (indoor) | UE (indoor) | * | * | 100 | 98,5 | 4 |
|  | UE (outdoor) | * | * | 70 | 37,7 | 6 |
|  | eNB (outdoor) | 5,5 | 0,8 | 85 | 54,11 | N/A |
| MBANS home (outdoor) | UE (indoor) | * | * | Covered | Covered | 7 |
|  | UE (outdoor) | * | * | 83 | 52,6 | 5 |
|  | eNB (outdoor) | 18,5 | 3,7 | 99 | 76,6 | 2 |

All these results need correction to I/N=-6dB

* If the LTE system is modeled as OFDMA, and UE is the victim, we cannot control where UE will be deployed and/or how far MBANs will be dropped from UE. Simulating the LTE system as generic, we can obtain the probability of a UE being interfered when a MBAN is transmitting close to a UE (i.e. 6-20m).

** No simulations because MBANs home(indoor)13dBm to UE(outdoor) is worse case.

Conclusion:

The MCL calculations and the SEAMCAT simulations show interference in both direction between MBANS and LTE, for both co-channel and adjacent channel cases. The co-channel cases would be the most difficult to handle. The impact of interference from MBANS on the LTE is more severe if MBANS are permitted outside of the healthcare facilities. Therefore, it is the conclusion that MBANS must comply with the Commission decision of 9 November 2006 on harmonization of the radio spectrum for use by short-range devices. SRDs in general can share frequency bands with other services only if they are not causing harmful interference, and furthermore no protection against interference can be guaranteed to users of SRDs. It is the responsibility of manufacturers of SRDs to protect such devices against harmful interference from radio communications services.

Mitigation techniques such as an improved MBAN Spectrum Emission Mask would mitigate the impact of MBANS on LTE system

/*END OF KEY PARTS OF ERICSSON CONTRIBUTION*/

[The analysis in Section 4.3.1.1 unveils a high interference potential from LTD DL into MBANSs.{to be revised which cases to mention here} Hence in order for both systems to co-exist, mitigation measures are necessary. Such mitigation measures may be based on geographical considerations, e.g. by selecting a different frequency band for LTE near healthcare facilities or by having MBANSs select a portion of the 2360 MHz-2400 MHz spectrum that is not used by LTE in the vicinity of the healthcare facility.

The aforementioned mitigation measures are not regarded as feasible for MBANSs operation in home and in ambulance. Hence in-band compatibility between LTE and home/ambulance MBANSs does not appear possible in the 2360—2400 MHz frequency range.]

[when the conclusions of this section are revisited, it should be considered to use the structure of other section (i.e) differentiation between MBANS types]

## SAP/SAB (ENG/OB) video links

Services Ancillary to Programme making / Services Ancillary to Broadcasting (SAP/SAB) are also known as Electronic News Gathering / Outside Broadcasting (ENG/OB). According to ECC Report 002 [ ], typical SAP/SAB applications in this frequency range are the video reportage links, with signal being transmitted to transportable studio from handheld or vehicle mounted cameras. Sometimes such links might be also used in airborne configuration, e.g. with the reportage camera being mounted on board a helicopter. However in this case the receiver is still a ground-based studio.

The parameters of SAP/SAB links as victim receivers used in this study were mainly taken from ECC Report 100 [ ], which considers interference to similar SAP/SAB links in the 3400 – 3600 MHz band. ECC Report 02 [], ECC Report 38 [] and ECC Report 02, and ETSI EN 300 744 [] were also consulted.

Two different types of SAP/SAB video links are considered, the parameters of which are summarized in the following tables.

Table 38: SAP/SAB video link parameters (Type 1 - mobile link, airborne)

|  | Parameter | Value |
| --- | --- | --- |
| Airborne Transmitter | Bandwidth (MHz) | 8 |
|  | Max Tx power (dBm) | 30 |
|  | Antenna gain (dBi) | 4 (omnidirectional) |
|  | Antenna height (m) | Up to 700. (375 assumed) |
| Ground Receiver (on vehicle) | Receiver sensitivity (dBm) | -89.8 |
|  | Antenna gain (dBi) | 10 (omnidirectional)[ |
|  | Antenna height (m) | 3 |
|  | C/I objective (dB) | 16 |

Table 39: SAP/SAB video link parameters (Type 2 - terrestrial link)

|  | Parameter | Value |
| --- | --- | --- |
| Transmitter | Bandwidth (MHz) | 8 |
|  | Max Tx power (dBm) | 43 |
|  | Antenna gain (dBi) | 0 (omnidirectional) |
|  | Antenna height (m) | 1.5 |
| Receiver | Receiver sensitivity (dBm) | -99 |
|  | Antenna gain (dBi) | 17 (directional) |
|  | Antenna diagram | See Figure 29 |
|  | Antenna height (m) | 1.5 |
|  | C/I objective (dB) | 10 |

Figure 29: Antenna radiation diagram of SAP/SAB receiver (Type 2)

According to ECC Report 38, SAP/SAB receivers on the roof of communications vehicles can use both omnidirectional and directional antennas. For ‘type 1’ video links, omnidirectional antennas have been assumed as worst case scenario. The use of directional antennas would lead to significantly lower levels of interference from MBANSs to SAP/SAB and vice versa.

### Healthcare Facility MBANSs

Table 36: MCL calculation in 2360 – 2400 MHz band between healthcare facility MBANSs and SAP/SAB video links

| Victim characteristics | Units | SAP/SAB (type 1) | MBANSs | SAP/SAB (type 2) | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 8 | 3 | 8 | 3 |
| Receiver noise figure | dB | 5 | 10 | 2 | 10 |
| Receiver antenna height | m | 3 | 1.5 | 1.5 | 1,5 |
| Receiver antenna gain | dBi | 10 | 0 | 17 | 0 |
| Operating frequency | MHz | 2380 | 2380 | 2380 | 2380 |
| N, receiver thermal noise | dBm | -99.8 | -99.1 | -102.8 | -99.1 |
| I/N objective | dB | -6 | 0 | -6 | 0 |
| Interferer’s characteristics |  | MBANSs | SAP/SAB (type 1) | MBANSs | SAP/SAB (type 2) |
| e.i.r.p | dBm | 0 | 34 | 0 | 43 |
| Bandwidth | MHz | 3 | 8 | 3 | 8000 |
| BW correction factor | dB | 0 | -4.26 | 0 | -4.26 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation | dB | 13 | 13 | 13 | 13 |
| Antenna height | m | 1,5 | 375 | 1,5 | 1.5 |
| Minimum path loss | dB | 102.8 | 115.8 | 112.8 | 124.8 |
| Interference distance FSL model (propagation exponent n=2) | km | 1.38 | 6.19 | 4.38 | 17.43 |
| Interference distance FSL model (propagation exponent n=3) | km | 0.12 | 0.34 | 0.27 | 0.67 |

#### Interference from MBANSs to SAP/SAB video links

The considered interference scenario is depicted in Figure 31 and has been simulated in SEAMCAT. The simulation settings and results are summarized in Table 39. The interference probability from MBANSs into SAP/SAB is low.

Figure 30: Interference scenario – healthcare facility MBANS into SAP/SAB video link (type1)

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Victim SAP/SAB receiver

Wanted signal paths

Interference paths

SAP/SAB transmitter

Table 37: Interference from healthcare facility MBANS to SAP/SAB - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): | SAP/SAB (type 1) | SAP/SAB (type 2) |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 8 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 10 Km) | Extended Hata, urban, outdoor→outdoor, below roof (user-defined radius, 500 m) |
| Interfering Link (ILK): | MBANS |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm | 0 dBm (10% probability) |
| ILT density | 40/km2 |  |
| ILT probability of transmission | 0.1 | 1 |
| ILT → VLR number of active Tx | - | 31 |
| ILT → VLR simulation radius | - | 500 m |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, below roof |  |
| ILT → VLR positioning mode | Closest interferer | None |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -71.66 (4.35) | -125.25 (17.95) |
| iRSSunwanted, dBm (Std.dev., dB) | -143.3 (19.3) | -144.17 (38.22) |
| Probability of interference (%) (C/I = 16 dB, type 1) (C/I = 10 dB, type 2) | 1.6 | 3.8 |

#### Interference from SAP/SAB video links to MBANSs

The considered interference scenario is depicted in Figure 32 and has been simulated in SEAMCAT. The simulation settings and results are summarized in Table 40. The interference probability from SAP/SAB into MBANSs is very low.

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Wanted SAP/SAB receiver

Wanted signal paths

Interference paths

SAP/SAB transmitter

Figure 31: Interference scenario – SAP/SAB video link (type 1) into healthcare facility MBANS

Table 38: Interference from SAP/SAB to healthcare facility MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): | MBANS |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |
| Interfering Link (ILK): | SAP/SAB (type 1) | SAP/SAB (type 2) |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 8 MHz |  |
| ILT Tx power | 30 dBm | 43 dBm |
| ILT density | 1/km2 |  |
| ILT probability of transmission | 1 | 1 |
| ILT → VLR number of active Tx | 1 | 1 |
| ILT → VLR simulation radius | 564 m | 500 m |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof | Extended Hata, urban, outdoor→indoor, below roof |
| ILT → VLR positioning mode | Uniform density | None |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | 45.17 (4.47) | -45.16 (4.43) |
| iRSSunwanted, dBm (Std.dev., dB) | -93.52 (12.94) | -119.88 (18.44) |
| Probability of interference (%) (C/I = 15 dB) | 0.9 | 1.0 |

### Home MBANSs

Table 39: MCL calculation in 2360 – 2400 MHz band between home 
MBANSs and SAP/SAB video links

| Victim characteristics | Units | SAP/SAB (type 1) | MBANSs | SAP/SAB (type 2) | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 8 | 3 | 8 | 3 |
| Receiver noise figure | dB | 5 | 10 | 2 | 10 |
| Receiver antenna height | m | 3 | 1,5 | 1.5 | 1.5 |
| Receiver antenna gain | dBi | 10 | 0 | 17 | 0 |
| Operating frequency | MHz | 2380 | 2380 | 2380 | 2380 |
| N, receiver thermal noise | dBm | -99,8 | -99,1 | -102.8 | -99.1 |
| I/N objective | dB | -6 | 0 | -6 | 0 |
| Interferer’s characteristics |  | MBANSs | SAP/SAB (type 1) | MBANSs | SAP/SAB (type 2) |
| e.i.r.p | dBm | 13 | 34 | 13 | 43 |
| Bandwidth | MHz | 3 | 8 | 3 | 8 |
| BW correction factor | dB | 0 | -4,26 | 0 | -4.26 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation (Note 1) | dB | 0 (outdoor) | 0 (outdoor) | 0 (outdoor) | 0 (outdoor) |
| Antenna height | m | 1,5 | 375 | 1.5 | 1.5 |
| Minimum path loss | dB | 128,8 | 128,8 | 138.8 | 137.8 |
| Interference distance FSL model (propagation exponent n=2) | km | 27.63 | 27.63 | 87.37 | 77.87 |
| Interference distance FSL model (propagation exponent n=3) | km | 0.91 | 0.91 | 1.97 | 1.82 |

Note 1: Worst case configuration assumed, in which MBANS transmitter or receiver is outdoors.

#### Interference from MBANSs to SAP/SAB video links

Wanted signal paths

Interference paths

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Home

13 dB building penetration loss

Victim SAP/SAB receiver

SAP/SAB transmitter

Figure 32: Interference scenario – healthcare facility MBANS into SAP/SAB video link (type 1)

Table 40: Interference from home MBANS to SAP/SAB - Settings and results

| Simulation input/output parameters | Settings / Results |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Victim Link (VLK): | SAP/SAB (type 1) | SAP/SAB (type 2) |  |  |  |
| VLK frequency | 2380 MHz |  |  |  |  |
| VLK bandwidth | 8 MHz |  |  |  |  |
| VLT → VLR path | FSL (user-defined radius, 10 Km) | Extended Hata, outdoor→outdoor, below roof (user-defined radius, 500 m) |  |  |  |
|  |  | urban | suburban | rural |  |
| Interfering Link (ILK): | MBANS |  |  |  |  |
| ILK frequency | 2380 MHz |  |  |  |  |
| ILK bandwidth | 3 MHz |  |  |  |  |
| ILT Tx power | 13 dBm | 13 dBm (2% probability) |  |  |  |
| ILT density | 10/km2 |  |  |  |  |
| ILT probability of transmission | 0.02 | 1 |  |  |  |
| ILT → VLR number of active Tx | 1 | 8 |  |  |  |
| ILT → VLR simulation radius | 1.26 km | 500 m |  |  |  |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, below roof | Extended Hata, indoor→outdoor, below roof |  |  |  |
|  |  | urban | suburban |  | rural |
| ILT → VLR positioning mode | Uniform density | None |  |  |  |
| Simulation results |  |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -71.66 (4.39) | -125.57 | -125.47 | -125.51 |  |
| iRSSunwanted, dBm (Std.dev., dB) | -144.3 (13.9) | -293.93 | -283.09 | -264.36 |  |
| Probability of interference (%) (C/I = 16 dB, type 1) (C/I = 10 dB, type 2) | 0.4 | 0.5 | 1.4 |  | 3.9 |

#### Interference from SAP/SAB video links to MBANSs

Wanted signal paths

Interference paths

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Home

13 dB building penetration loss

Wanted SAP/SAB receiver

SAP/SAB transmitter

Figure 33: Interference scenario – SAP/SAB video link (type 1) into home MBANS

Table 41: Interference from SAP/SAB to home MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): | MBANS |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 10 m) |  |
| Interfering Link (ILK): | SAP/SAB (type 1) | SAP/SAB (type 2) |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 8 MHz |  |
| ILT Tx power | 30 dBm | 43 dBm |
| ILT density | 1/km2 |  |
| ILT probability of transmission | 1 |  |
| ILT → VLR number of active Tx | 1 |  |
| ILT → VLR simulation radius | 564 | 500 |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof | Extended Hata, urban, outdoor→indoor, below roof |
| ILT → VLR positioning mode | Uniform density | None |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -42.62 (4.48) | -42.62 (4.45) |
| iRSSunwanted, dBm (Std.dev., dB) | -93.58 (12.95) | -119.77 (18.41) |
| Probability of interference (%) (C/I = 15 dB) | 0.6 | 1.0 |

### Ambulance MBANS

Table 42: MCL calculation in 2360 – 2400 MHz band between ambulance MBANSs and SAP/SAB video links

| Victim characteristics | Units | SAP/SAB (type 1) | MBANSs | SAP/SAB (type 2) | MBANSs |
| --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 8 | 3 | 8 | 3 |
| Receiver noise figure | dB | 5 | 10 | 2 | 10 |
| Receiver antenna height | m | 3 | 1.5 | 1.5 | 1.5 |
| Receiver antenna gain | dBi | 10 | 0 | 17 | 0 |
| Operating frequency | MHz | 2380 | 2380 | 2380 | 2380 |
| N, receiver thermal noise | dBm | -99.8 | -99.1 | -102.8 | -99.1 |
| I/N objective | dB | -6 | 0 | -6 | 0 |
| Interferer’s characteristics |  | MBANSs | SAP/SAB (type 1) | MBANSs | SAP/SAB (type 2) |
| e.i.r.p | dBm | 0 | 34 | 0 | 43 |
| Bandwidth | MHz | 3 | 8 | 3 | 8 |
| BW correction factor | dB | 0 | -4.26 | 0 | 4.26 |
| NFD (adjacent band interf) | dB | 0 | 0 | 0 | 0 |
| Wall attenuation | dB | 7 | 7 | 7 | 7 |
| Antenna height | m | 1.5 | 375 | 1.5 | 1.5 |
| Minimum path loss | dB | 108.8 | 121.8 | 118.8 | 130.8 |
| Interference distance FSL model (propagation exponent n=2) | km | 2.76 | 12.34 | 8.74 | 34.78 |
| Interference distance FSL model (propagation exponent n=3) | km | 0.20 | 0.53 | 0.42 | 1.07 |

#### Interference from MBANSs to SAP/SAB video links

Wanted signal paths

Interference paths

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB penetration loss

Victim SAP/SAB receiver

SAP/SAB transmitter

Figure 34: Interference scenario – ambulance MBANS into SAP/SAB video link (type 1)

Table 43: Interference from ambulance MBANSs to SAP/SAB - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): | SAP/SAB (type 1) | SAP/SAB (type 2) |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 8 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 10 Km) | Extended Hata, urban, outdoor→outdoor, below roof (user-defined radius, 500 m) |
| Interfering Link (ILK): | MBANS |  |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm | 0 dBm (10% probability) |
| ILT density | 5/km2 |  |
| ILT probability of transmission | 0.1 | 1 |
| ILT → VLR number of active Tx | 2 | 4 |
| ILT → VLR simulation radius | 1.13 | 500 m |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, below roof |  |
| ILT → VLR positioning mode | Uniform density | None |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -71.69 (4.34) | -125.53 (17.96) |
| iRSSunwanted, dBm (Std.dev., dB) | -141.02 (13.74) | -264.05 (86.76) |
| Probability of interference (%) (C/I = 16 dB, type 1) (C/I = 10 dB, type 2) | 0.7 | 0.7 |

#### Interference from SAP/SAB video links to MBANSs

Wanted signal paths

Interference paths

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB penetration loss

Victim SAP/SAB receiver

SAP/SAB transmitter

Figure 35: Interference scenario – SAP/SAB video link (type 1) into healthcare facility MBANS

Table 44: Interference from SAP/SAB to ambulance MBANSs - Settings and result

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): | MBANS |  |
| VLK frequency | 2380 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |
| Interfering Link (ILK): | SAP/SAB (type 1) | SAP/SAB (type 2) |
| ILK frequency | 2380 MHz |  |
| ILK bandwidth | 8 MHz |  |
| ILT Tx power | 30 dBm | 43 dBm |
| ILT density | 1/km2 |  |
| ILT probability of transmission | 1 |  |
| ILT → VLR number of active Tx | 1 |  |
| ILT → VLR simulation radius | 564 m | 500 m |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof | Extended Hata, urban, outdoor→indoor, below roof |
| ILT → VLR positioning mode | Uniform density | None |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -45.22 (4.46) | -45.19 (4.47) |
| iRSSunwanted, dBm (Std.dev., dB) | -87.73 (12.87) | -113.89 (18.57) |
| Probability of interference (%) (C/I = 15 dB) | 2.6 | 1.5 |

### Summary SAP/SAB

Healthcare facility MBANSs:

Average case simulation results suggest coexistence between SAP/SAB and healthcare facility MBANSs. Even in unfavorable conditions, the required distances between SAP/SAB and healthcare facilities are in a range that corresponds to most situations in practice. Therefore coexistence with MBANSs within healthcare facilities appears possible.

Home MBANSs:

Average case simulation results suggest good coexistence between SAP/SAB and home MBANSs. However in unfavorable conditions, the required distances between SAP/SAB and the homes of MBANS users would be unpractical. Such potential interference to SAP/SAB applications—even with low deployment values of 10 MBANSs per km2—lead to the conclusion that coexistence is not possible in areas where SAP/SAB is used in the 2300 — 2400 MHz band.

Ambulance MBANSs:

Average case simulation results suggest good coexistence between SAP/SAB and ambulance MBANSs. However in unfavorable conditions, the required distances between SAP/SAB and ambulances could not be guaranteed, due to the high mobility of the latter. The potential interference to SAP/SAB applications—even with low deployment values of 5 MBANSs per km2—lead to the conclusion that coexistence is not possible in areas where SAP/SAB is used in the 2300 — 2400 MHz band.

## ADJACENT Wideband data transmission systems

This section studies the compatibility between MBANSs operating in the 2300 — 2400 MHz range and wideband data transmission systems–represented by WLAN– operating in the adjacent 2400 — 2483.5 MHz range.  Due to the equivalence to the co-channel compatibility study in Section 5.1, this section is presented in compact form. More details and considerations can be found in Section 5.1.

Based on the WLAN standard [] the last channel of wideband data transmission on the lower side of the band ends at 2402 MHz. That results in 2 MHz buffer between WLAN and the possible MBANs implementation in the band studied in this section (2300 — 2400 MHz). However, it was considered appropriate to also study possible interference between these technologies in these two adjacent bands. The MCL calculations for 20 MHz WLAN channel bandwidth are presented in Table 45. The same approach as in the co-channel compatibility analysis in Section 5.1 was used.

Table 45: Adjacent MCL calculation in 2400 MHz band limits between MBANSs and WLAN

| Victim characteristics | Units | WLAN 802.11 AP | WLAN 802.11 AP | MBANSs |
| --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 20 | 20 | 3 |
| Receiver noise figure | dB | 10 | 10 | 10 |
| Receiver antenna height | m | 3 | 3 | 1.5 |
| Receiver antenna gain | dBi | 5 | 5 | 0 |
| Operating frequency | MHz | 2412 | 2412 | 2398.5 |
| N, receiver thermal noise | dBm | -90.82 | -90.82 | -99.1 |
| I/N objective | dB | 0 | 0 | 0 |
| Interferer’s characteristics |  | MBANSs (healthcare facility) | MBANSs (home) | WLAN 802.11 AP |
| e.i.r.p | dBm | 0 | 13 | 17 |
| Bandwidth | MHz | 3 | 3 | 20 |
| BW correction factor | dB | 0 | 0 | -8.24 |
| NFD (adjacent band interf.) | dB | -25 | -25 | -40 |
| Antenna height | m | 1.5 | 1.5 | 3 |
| Minimum path loss | dB | 70.8 | 67.8 | 67.8 |
| Interference distance FSL model | km | 0.03 | 0.02 | 0.02 |

### Healthcare facility MBANSs

The SEAMCAT simulations are summarized in the following table. The study approach is identical to that Section 5.1.1. The parameter and scenario choices are described in that section of the report. For obtaining the interference probability in SEAMCAT, the combination of unwanted and blocking response was used. A value of 30 dB blocking response was used for MBANSs.

Table 46: Interference in 2400 MHz band limits between healthcare facility MBANSs and WLAN – Simulation results

| Interferer | Victim | Interferer central frequency (MHz) | Victim central frequency (MHz) | Simulation parameters | Interference probability / bitrate degradation |
| --- | --- | --- | --- | --- | --- |
| MBANS | WLAN (20 MHz) | 2398.5 | 2412 | See Section 5.1.1.1 | 0.1 % |
|  | WLAN (40 MHz) |  | 2422.5nz) approach is identical to that Section llowing table. gies in these two adjacent bands. |  | 0.1 % |
| WLAN (20 MHz) | MBANS | 2412 | 2398.5 | See Section 5.1.1.2 | 4 % |
| WLAN (40 MHz) |  | 2422 |  |  | 3.7 % |

### Home MBANSs

The SEAMCAT simulations are summarized in the following table. The study approach is identical to that Section 5.1.1. The parameter and scenario choices are described in that section of the report. For obtaining the interference probability in SEAMCAT, the combination of unwanted and blocking response was used. A value of 30 dB blocking response was used for MBANSs.

Table 47: Interference in 2400 MHz band limits between home MBANSs and WLAN – Simulation results

| Interferer | Victim | Interferer central frequency (MHz) | Victim central frequency (MHz) | Simulation parameters | Interference probability / bitrate degradation |
| --- | --- | --- | --- | --- | --- |
| MBANS | WLAN (20 MHz) | 2398.5 | 2412 | See Section 5.1.2.1 | 0 % |
|  | WLAN (40 MHz) |  | 2422.5nz) approach is identical to that Section llowing table. gies in these two adjacent bands. |  | 0.1 % |
| WLAN (20 MHz) | MBANS | 2412 | 2398.5 | See Section 5.1.2.2 | 3.1 % |
| WLAN (40 MHz) |  | 2422 |  |  | 3 % |

### Summary ADJACENT wideband data transmission systems

[

All simulation-based average-case analyses yielded 0% to 0.1% probability of interference from MBANSs. In the other coexistence direction, the obtained results range from 3% to 4% probability of interference from wideband data transmission systems.

Wideband data transmission systems operating in the 2400 — 2483.5 MHz range and MBANSs in healthcare facilities, homes, and ambulances operating in the 2300 — 2400 MHz range are compatible.

]

## ADJACENT Deep Space Research Service (SRS)

The adjacent band 2290 − 2300 MHz is allocated for deep space research. Specifically, it is for the downlink from space satellite to Earth stations. Since the satellite signal from space is weak, the interference from deep space satellite to MBANSs is considered negligible. Therefore, only the interference from MBANSs to deep space Earth station receivers is considered. According to Recommendation ITU-R SA.1157 [xx], a commonly acceptable protection criterion for deep space research Earth station is -222 dBW/Hz at the input of the Earth station receiver, which is equivalent to -132 dBm/MHz.

To receive the weak satellite signal from space, Earth station receivers use a high-gain directional antenna that points to satellite. A typical Earth station receive antenna pattern is shown in Figure xx and defined as below, where φ (in degrees) is the angle between the axis of main beam and the direction considered.

=60 dBi,  0 &≤φ<132-25logφdBi,  &1 ≤φ<48-10 dBi, 48 ≤φ≤180

Figure 36: Reference Earth station antenna pattern (from Recommendation ITU-R SA. 509)

Such antenna pattern comes from Recommendation ITU-R SA. 509 [xx] and is used in the following SEAMCAT simulations.  In its normal working mode, Earth stations point the receive antenna to the sky to track satellite in azimuth and in elevation, leaving MBANS signals outside its main beam. Usually, φ, the angle between the axis of main beam and the arriving direction of MBANS signal, is larger than 10o. In the MCL calculations, φ is assumed to be 5o to model a worst-case scenario. That means that the assumed receive antenna gain is 14.5 dBi.

In the studied scenarios, it is assumed that both MBANs and an Earth station receiver operate at the 2300 MHz boundary and their operating channels are adjacent to each other. This means that there is no guard band considered here. With such assumption, a conservative NFD, -20 dB, is used in the MCL calculations.

### Healthcare facility MBANSs

The MCL calculation for the healthcare facility MBANS case is summarized in the following table.

Table 48: Adjacent MCL calculation in 2300 MHz band limits between healthcare facility MBANSs and Deep Space SRS

| Victim characteristics | Units | Deep Space SRS Station |  |
| --- | --- | --- | --- |
| Receiver antenna gain for the ILT → VLR path | dBi | 14.5 |
| Protection Criteria | dBm/MHz | -132 |
| Interferer’s characteristics |  | MBANSs |
| e.i.r.p | dBm | 0 |
| Operating frequency | MHz | 2301.5 |
| Bandwidth | MHz | 3 |
| NFD (adjacent band interf.) | dB | -20 |
| PSD in adjacent channel below 2300 MHz | dBm/MHz | -24.8 |
| Wall attenuation | dB | 13 |
| Minimum path loss | dB | 108.7 |
| Interference distance FSL model (propagation exponent n=2) | km | 2.82 |
| Interference distance FSL model (propagation exponent n=3) | km | 0.2 |

A more realistic case is studied with SEAMCAT, and the considered interference scenario is depicted in Figure 37. The SEAMCAT simulation settings and results are summarized in Table 49.

Figure 37: Interference scenario – healthcare facility MBANS into Deep Space SRS

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Victim Earth station receiver

Wanted signal paths

Interference paths

Table 49: Interference from healthcare facility MBANS to Deep Space SRS - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): Deep Space (Satellite-to-Earth station) |  |  |
| VLK frequency (Note 1) | 2299.5 MHz |  |
| VLK bandwidth (Note 1) | 1 MHz |  |
| VLT → VLR path (Note 1) | FSL (user-defined radius, 800 Km) |  |
| VLR dRSS (Note 1) | -100 dBm |  |
| VLR noise floor (Note 2) | -114 dBm |  |
| VLR receive antenna pattern | See Figure 36 |  |
| VLR receive antenna height | 20 m |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2301.5 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm |  |
| ILT density | 40/km2 |  |
| ILT probability of transmission | 0.1 |  |
| ILT → VLR interfering path | Extended Hata, suburban, indoor→outdoor, above roof |  |
| ILT → VLR positioning mode | Closest interferer |  |
| ILT → VLR protection distance | 0 | 1.5 km |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) (nominal) | -100 (0) (nominal) |
| iRSSunwanted, dBm (Std.dev., dB) | -157.9 (16.2) | -181.67 (11.31) |
| Probability of interference (%) (I/N = -18 dB) (Note 2) | 6.4 | 0.00 |

Note 1: Since the protection criterion used is -222 dBw/Hz, the VLK parameters are nominal and do not affect the final interference probability.

Note  2: The combination of  I/N = -18 dB and N = -114 dBm is used to represent the -132 dBm/MHz protection criterion.

In RECOMMENDATION ITU-R SA.1157, it is stated “…that calculation of interference that may result from atmospheric and precipitation effects be based on weather statistics that apply for 0.001% of the time”. Following ECC Report 172[], it is considered that this reference percentage of time is 0.001 % is equivalent to the probability of interference in SEAMCAT. Due to the limited resolution of SEAMCAT (0.01%), an interference probability result of 0.00% was used as reference to find the protection distance. The simulation results show that in order to achieve compatibility, a protection distance of 1.5 km around Earth station receivers is required. Even in the worst case studied in the MCL calculation, the required separation distance under the free space channel model assumption is only 2.82 km. Given the limited number of deep space research Earth station receivers in the CEPT area, it is considered feasible to impose a restriction on MBANSs to operate only within healthcare facilities that lie beyond a predefined protection distance (e.g. 1.5 km).

### Home MBANSs

The MCL calculation for the home MBANS case is summarized in the following table.

Table 50: Adjacent MCL calculation in 2300 MHz band limits between home MBANSs and Deep Space SRS

| Victim characteristics | Units | Deep Space SRS Station |  |
| --- | --- | --- | --- |
| Receiver antenna gain for the ILT → VLR path | dBi | 14.5 |
| Protection Criteria | dBm/MHz | -132 |
| Interferer’s characteristics |  | MBANSs |
| e.i.r.p | dBm | 13 |
| Operating frequency | MHz | 2301.5 |
| Bandwidth | MHz | 3 |
| NFD (adjacent band interf.) | dB | -20 |
| PSD in adjacent channel below 2300 MHz | dBm/MHz | -11.8 |
| Wall attenuation | dB | 0 (outdoor) |
| Minimum path loss | dB | 134.7 |
| Interference distance FSL model (propagation exponent n=2) | km | 56.2 |
| Interference distance FSL model (propagation exponent n=3) | km | 1.5 |

A more realistic case is studied with SEAMCAT, and the considered interference scenario is depicted in Figure 38. The SEAMCAT simulation settings and results are summarized in Table 51.

Figure 38: Interference scenario – home MBANS into Deep Space SRS

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Victim Earth station receiver

Wanted signal paths

Interference paths

Table 51: Interference from home MBANS to Deep Space SRS - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): Deep Space (Satellite-to-Earth station) |  |  |
| VLK frequency (Note 1) | 2299.5 MHz |  |
| VLK bandwidth (Note 1) | 1 MHz |  |
| VLT → VLR path (Note 1) | FSL (user-defined radius, 800 km) |  |
| VLR dRSS (Note 1) | -100 dBm |  |
| VLR noise floor (Note 2) | -114 dBm |  |
| VLR receive antenna pattern | See Figure 36 |  |
| VLR receive antenna height | 20 m |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2301.5 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 13 dBm |  |
| ILT density | 10/km2 |  |
| ILT probability of transmission | 0.02 |  |
| ILT → VLR interfering path | Extended Hata, suburban, indoor→outdoor, above roof |  |
| ILT → VLR positioning mode | Uniform density |  |
| ILT → VLR protection distance | 0 | 2.8 km |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) (nominal) | -100 (0) (nominal) |
| iRSSunwanted, dBm (Std.dev., dB) | -158.3 (3.9) | -178.81 (11.29) |
| Probability of interference (%) (I/N = -18 dB) (Note 2) | 4.2 | 0.00 |

Note 1: Since the protection criterion used is -222 dBw/Hz, the VLK parameters are nominal and do not affect the final interference probability.

Note  2: The combination of  I/N = -18 dB and N = -114 dBm is used to represent the -132 dBm/MHz protection criterion.

The simulation results show that in order to achieve compatibility, a protection distance of 2.8 km around Earth station receivers is required. In the worst case studied in the MCL calculation, the required separation distance under the free space channel model assumption can go beyond 50 km. Therefore, and also considering the potentially widespread deployment of patient houses, it is concluded that the coexistence between home MBANS and deep space research in this adjacent channel case cannot be guaranteed with a protection distance.

### Ambulance MBANSs

The MCL calculation for the ambulance MBAN case is summarized in the following table.

Table 52: Adjacent MCL calculation in 2300 MHz band limits between ambulance MBANSs and Deep Space SRS

| Victim characteristics | Units | Deep Space SRS Station |  |
| --- | --- | --- | --- |
| Receiver antenna gain for the ILT → VLR path | dBi | 14.5 |
| Protection Criteria | dBm/MHz | -132 |
| Interferer’s characteristics |  | MBANSs |
| e.i.r.p | dBm | 0 |
| Operating frequency | MHz | 2301.5 |
| Bandwidth | MHz | 3 |
| NFD (adjacent band interf.) | dB | -20 |
| PSD in adjacent channel below 2300 MHz | dBm/MHz | -24.8 |
| Wall attenuation | dB | 7 |
| Minimum path loss | dB | 114.7 |
| Interference distance FSL model (propagation exponent n=2) | km | 5.6 |
| Interference distance FSL model (propagation exponent n=3) | km | 0.3 |

A more realistic case is studied with SEAMCAT, and the considered interference scenario is depicted in Figure 39. The SEAMCAT simulation settings and results are summarized in Table 53.

Figure 39: Interference scenario – ambulance MBANS into Deep Space SRS

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB  penetration loss

Wanted signal paths

Interference paths

Victim Earth station receiver

Table 53: Interference from ambulance MBANS to Deep Space SRS - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): Deep Space (Satellite-to-Earth station) |  |  |
| VLK frequency (Note 1) | 2299.5 MHz |  |
| VLK bandwidth (Note 1) | 1 MHz |  |
| VLT → VLR path (Note 1) | FSL (user-defined radius, 800 km) |  |
| VLR dRSS (Note 1) | -100 dBm |  |
| VLR noise floor (Note 2) | -114 dBm |  |
| VLR receive antenna pattern | See Figure 36 |  |
| VLR receive antenna height | 20 m |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2301.5 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm |  |
| ILT density | 5/km2 |  |
| ILT probability of transmission | 0.1 |  |
| ILT → VLR interfering path | Extended Hata, suburban, indoor→outdoor, above roof |  |
| ILT → VLR positioning mode | Uniform density |  |
| ILT → VLR protection distance | 0 | 1.9 km |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) (nominal) | -100 (0) (nominal) |
| iRSSunwanted, dBm (Std.dev., dB) | -158.2 (14.29) | -179.9 (11.29) |
| Probability of interference (%) (I/N = -18 dB) (Note 2) | 4.4 | 0.00 |

Note 1: Since the protection criterion used is -222 dBw/Hz, the VLK parameters are nominal and do not affect the final interference probability.

Note  2: The combination of  I/N = -18 dB and N = -114 dBm is used to represent the -132 dBm/MHz protection criterion.

The simulation results show that in order to achieve compatibility, a protection distance of 1.9 km around Earth station receivers is required. Considering the mobility of ambulance MBANSs, it is concluded that the coexistence between ambulance MBANSs and deep space research in this adjacent channel case cannot be guaranteed.

### Summary Deep Space SRS

[

Healthcare facility MBANSs:

According to the results presented above, in worst-case situations—modelled with MCL calculations—the interference distance between Deep Space SRS and MBANSs is 0.2 km - 2.8 km (MBANS as interferer). The simulation-based average-case analysis yielded 6.4% interference probability. The result improves significantly with a separation distance of 1.5 km, yielding 0.001% interference probability, as required to protect Deep Space SRS.

Deep space SRS has strict interference probability requirements. Nonetheless, compatibility between healthcare-facility MBANSs and Deep Space SRS is feasible with a reasonable geographical separation distance (e.g. 1.5 km) between the healthcare facility using MBANSs and the SRS Earth station. The number of SRS Earth stations is limited and their location is known in many countries. Furthermore, since healthcare facility MBANS deployments may have better characteristics in reality than those assumed in the studies (e.g. better out of band emission performance, smaller receive antenna gain for the ILT->VLR link, guard band, etc.), no adjacent band compatibility issues are expected with healthcare facility MBANSs.

Home MBANSs and ambulance MBANSs:

In worst-case situations—modelled with MCL calculations—the interference distance between Deep Space SRS and MBANSs is 1.5 km - 56 km (home MBANS as interferer) and 0.3 km – 5.6 km (ambulance MBANS as interferer). The simulation-based average-case analysis yielded a separation distance requirement of 2.8 km (home MBANSs) and 1.9 km (ambulance MBANSs).

Adjacent band compatibility between home MBANSs or ambulance MBANSs and Deep Space SRS may be possible in practice but cannot be guaranteed due to the widespread/mobile characteristics of home and ambulance MBANSs.]

# Compatibility Study for the Band 2400 – 2483.5 MHz

The compatibility studies presented in this section are structured in the following way:

- Section 5.x:  Introduction of the radio service / application analysed in that subsection, including all relevant technical parameters for the studies

- Section 5.x.1: Compatibility with healthcare facility MBANSs, including MCL calculations and SEAMCAT simulations

- Section 5.x.2: Compatibility with home MBANSs, including MCL calculations and SEAMCAT simulations

- Section 5.x.3: Compatibility with ambulance MBANSs mode, including MCL calculations and SEAMCAT simulations

- Section 5.x.4: Summary and conclusions for the analysed radio service / application

## Wideband data transmission systems

WLAN systems based on the IEEE 802.11 b/g standard have been selected as the most representative system for co-existence analysis with MBANS applications. Such systems have extremely proliferated both in residential and hospital areas, having achieved a large install base of Access Points (AP) for wireless data connectivity—e.g. enabling medical personnel to use their smartphones or similar devices for wirelessly accessing the hospital databases, localising medical equipment, download-uploading patients journal’s data, establishing VoIP communications, etc.

For these MCL studies, the typical WLAN was considered to operate with 20 MHz channel bandwidth (at -20dB) and 17 dBm e.i.r.p. Note also that because in the worst-case scenarios MBANS and WLAN Access Point (AP) are located inside the same building, direct LOS coupling should be assumed, suggesting the use of the Free Space Model for path loss calculations. The results of calculated MCL interference distances for the identified critical WLAN applications in the band 2400 – 2483.5 MHz are presented in Table 45 below. The interference distances obtained with the MCL method made it clear that a more detailed analysis with the SEAMCAT simulation tool was necessary.

Table 45: MCL calculation in 2400 – 2483,5 MHz band between MBANSs and WLAN

| Victim characteristics | Units | WLAN 802.11 AP | WLAN 802.11 AP | MBANSs |
| --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 20 | 20 | 3 |
| Receiver noise figure | dB | 10 | 10 | 10 |
| Receiver antenna height | m | 3 | 3 | 1.5 |
| Receiver antenna gain | dBi | 5 | 5 | 0 |
| Operating frequency | MHz | 2420 | 2420 | 2420 |
| N, receiver thermal noise | dBm | -90.82 | -90.82 | -99.1 |
| I/N objective | dB | 0 | 0 | 0 |
| Interferer’s characteristics |  | MBANSs (healthcare facility) | MBANSs (home) | WLAN 802.11 AP |
| e.i.r.p | dBm | 0 | 13 | 17 |
| Bandwidth | MHz | 1 – 5 | 1 – 5 1 | 20 |
| BW correction factor | dB | 0 | 0 | -8.24 |
| NFD (adjacent band interf.) | dB | 0 | 0 | 0 |
| Antenna height | m | 1.5 | 1.5 | 3 |
| Minimum path loss | dB | 95.8 | 108.8 | 107.8 |
| Interference distance FSL model | km | 0.61 | 2.72 | 2.43 |

Table 46: Wideband data transmission systems (RLAN) parameters

|  | Parameter | Value |  |
| --- | --- | --- | --- |
| RLAN Transmitter (AP) | Bandwidth (MHz) | 20 (IEEE 802.11b/g) | 40 (IEEE 802.11n) |
|  | Max Tx power (dBm) | 17 |  |
|  | Antenna gain (dBi) | 5 (omnidirectional assumed) |  |
|  | Antenna height (m) | 3 |  |
| RLAN Receiver | Receiver sensitivity (dBm) | -82 | -79 |
|  | Antenna gain (dBi) | 0 (omnidirectional) |  |
|  | Antenna height (m) | 1.5 |  |
|  | C/I objective (dB) | 10 |  |

Figure 40: RLAN emission masks. Left: IEEE 802.11g (20 MHz). Right: IEEE 802.11n (40 MHz)

### Healthcare facility MBANSs

#### Interference from MBANSs to wideband data transmission systems

Table 47: Interference from healthcare facility MBANSs into RLAN- Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): RLAN |  |  |
| VLK frequency | 2420 MHz |  |
| VLK bandwidth | 20 MHz | 40 MHz |
| VLT → VLR path | Extended Hata, urban, indoor→indoor, below roof (user-defined radius, 20 m) |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2420 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm (10% probability) |  |
| ILT density | 40/km2 |  |
| ILT probability of transmission | 1 |  |
| ILT → VLR interfering path | Extended Hata, urban, indoor→indoor, below roof |  |
| ILT → VLR positioning mode | Uniform density |  |
| ILT → VLR number of transmitters | 1 (for 90 m simulation radius) |  |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -53.75 (14) |  |
| iRSSunwanted, dBm (Std.dev., dB) | -304.08 (64.99) |  |
| Probability of interference (%) (C/I = 10 dB) | 0.2 |  |

#### Interference from wideband data transmission systems to MBANSs

Table 48: Interference from RLAN into healthcare facility MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |
| VLK frequency | 2420 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |
| Interfering Link (ILK): RLAN |  |  |
| ILK frequency | 2420 MHz |  |
| ILK bandwidth | 20 MHz | 40 MHz |
| ILT Tx power | 17 dBm (60% probability) |  |
| ILT density | 625/km2 |  |
| ILT probability of transmission | 1 |  |
| ILT → VLR interfering path | Extended Hata, urban, indoor→indoor, below roof |  |
| ILT → VLR positioning mode | Uniform density |  |
| ILT → VLR number of transmitters | 16 (for 90 m simulation radius) |  |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -45.3 (4.42) | -45.35 (4.43) |
| iRSSunwanted, dBm (Std.dev., dB) | -69.68 (13.99) | -72.48 (14.21) |
| Probability of interference (%) (C/I = 15 dB) | 23.2 | 19.0 |

### Home MBANSs

#### Interference from MBANSs to wideband data transmission systems

Table 49: Interference from home MBANSs into RLAN- Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): RLAN |  |  |
| VLK frequency | 2420 MHz |  |
| VLK bandwidth | 20 MHz | 40 MHz |
| VLT → VLR path | Extended Hata, urban, indoor→indoor, below roof (user-defined radius, 20 m) |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2420 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 13 dBm (2% probability) |  |
| ILT density | 10/km2 |  |
| ILT probability of transmission | 1 |  |
| ILT → VLR interfering path | Extended Hata, urban, indoor→indoor, below roof |  |
| ILT → VLR positioning mode | Uniform density |  |
| ILT → VLR number of transmitters | 1 (for 180 m simulation radius) |  |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -53.74 (14.15) |  |
| iRSSunwanted, dBm (Std.dev., dB) | -341 (38.7) |  |
| Probability of interference (%) (C/I = 10 dB) | 0.1 |  |

#### Interference from wideband data transmission systems to MBANSs

Table 50: Interference from RLAN into home MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |
| VLK frequency | 2420 MHz |  |
| VLK bandwidth | 3 MHz |  |
| VLT → VLR path | FSL (user-defined radius, 10 m) |  |
| Interfering Link (ILK): RLAN |  |  |
| ILK frequency | 2420 MHz |  |
| ILK bandwidth | 20 MHz | 40 MHz |
| ILT Tx power | 17 dBm (60% probability) |  |
| ILT density | 625/km2 |  |
| ILT probability of transmission | 1 |  |
| ILT → VLR interfering path | Extended Hata, urban, indoor→indoor, below roof |  |
| ILT → VLR positioning mode | Uniform density |  |
| ILT → VLR number of transmitters | 16 (for 90 m simulation radius) |  |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -42.72 (4.58) | -42.79 (4.49) |
| iRSSunwanted, dBm (Std.dev., dB) | -69.77 (14.11) | -72.92 (14.04) |
| Probability of interference (%) (C/I = 15 dB) | 18.8 | 14.5 |

### Ambulance MBANSs

The use of RLAN in ambulances is regarded as marginal, due to the mobility of ambulances and its detrimental effect on RLAN communications. Therefore it is considered unnecessary to analyse compatibility with ambulance MBANSs in further detail. The physical distance between in-building RLANs and in-ambulance MBANSs—together with the attenuation added by the building wall and ambulance chassis—provide enough evidence to expect good coexistence.

### Summary wideband data transmission systems

[

Healthcare facility MBANSs and home MBANSs:

According to the results presented above, in worst-case situations—modelled with MCL calculations—the interference distance between wideband data transmission systems and MBANSs is 0.6 km (healthcare facility MBANS as interferer),  2.7 km (home MBANS as interferer), and 2.4 km (MBANS as victim). The simulation-based average-case analysis yielded ≤0.2% interference probability from MBANSs into wideband data transmission systems. In the other coexistence direction, the obtained results lie in the range of 15% - 23% probability of interference from wideband data transmission systems.

Co-channel compatibility is not given between wideband data transmission systems and MBANSs in healthcare facilities and homes. Average case simulation results unveiled significant interference levels from wideband data transmission systems into MBANSs. Due to the widespread deployment of wideband data transmission systems (such as WLANs), separation-based mitigation measures are not practicable in healthcare facilities and patient homes.

Ambulance MBANSs:

Due to the absence or very low footprint of wideband data transmission systems within ambulances, co-channel compatibility with ambulance MBANSs can be expected.

]

## ISM

Some of the examples of ISM applications that could be considered as potential interferers to MBANSs for the hospital environment would be e.g. cauterising tools used as surgery aids in hospitals or microwave ovens, which may be found both in hospitals and patient homes. Regarding the cauterising tools, this study could not get hold of any reliable data that would allow deterministic characterising of RF emissions from these devices. Regarding microwave ovens, as discussed in previous ECC reports [ECC rep 149], it appears that the leakage of RF power from microwave ovens is very small (e.g. 0.2 mW/cm2 quoted in some sources [ ]) compared to other applications such as WLAN APs. Therefore for the above described reasons, both of these ISM devices were left out of this study.

## Non-specific SRDs

It was decided not to explicitly consider the compatibility between MBANSs and non-specific Short Range Devices (SRDs) beyond the considerations in Section 5.1, since WLAN systems are considered as the most critical representative of SRDs for co-existence analysis.

## Railway Applications

According to ETSI TR 101 557 [ ], MBANSs require 40 MHz of spectrum for their operation. It is hence proposed to limit MBANS operation to the 2400 – 2440 MHz band within the 2400 – 2483.5 MHz range. This limitation eliminates the need to analyse in-band coexistence with railway applications, which operate in the 2446 – 2454 MHz range [ERC/REC 70-03].

## RFID

As discussed in 5.4, it is proposed to limit MBANS operation to the 2400 – 2440 MHz band. This eliminates the need to analyse in-band coexistence with RFID applications, which operate in the 2446 – 2454 MHz range [ERC/REC 70-03].

## Amateur and Amateur Satellite

The range 2400 – 2450 MHz is allocated to both the amateur and amateur satellite service on a secondary basis. Terrestrial amateur service usage in this frequency range is most typically for multimedia and Amateur TV, with similar characteristics to the 2300 – 2400 MHz systems studied in Section 4.2. The only difference is the operation frequency, which does not have a significant influence on the interference probability results. Therefore the same results and conclusions as in Section 4.2 apply to amateur multimedia and amateur TV in the 2400 – 2450 MHz band. This section explicitly analyses compatibility with the amateur satellite service.

Recommendation ITU-R M.1732-1[] provides characteristics for stations operating in the amateur satellite service. They are summarized in Table 60.

Table 60: Characteristics of Amateur-Satellite systems in the space-to-Earth direction

| Parameter | CW-Morse | SSB, FM, Digital Voice, Data |
| --- | --- | --- |
| Transmitter Power (1) (dBW) | 10 | 10 |
| Transmitter Feeder Loss (dB) | 0.2 – 1 | 0.2 – 1 |
| Transmitting antenna gain (dBi). | 0 – 6 | 0 |
| Typical e.i.r.p.(dBW) | 0 – 15 | 9 – 15 |
| Antenna polarisation | Horizontal, Vertical, LHCP, RHCP | Horizontal, Vertical, LHCP, RHCP |
| Receiver IF bandwidth (kHz) | 0.4 | 2.7, 16, 50, 100 |
| Receiver Noise Figure (2) (dB) | 1 – 3 (typically: 1) | 1 – 3 (typically: 2) |

(1) Maximum power levels are determined by each administration.

(2) Receiver noise figures for bands above 50 MHz assume the use of low-noise preamplifiers.

In contrast to the assumptions in ITU M.1732-1[], most current amateur satellites are typically picosats (also called ‘cubesats’), which occupy slightly elliptical Sun-Synchronous low Earth orbits (LEO) of 600-800 km altitude. These smaller satellites have relatively low power and antenna gain. Whilst these are most typical, other scenarios can include reception of Digital Amateur TV downlinks from the International Space Station, and weak signal reception of longer range satellites. It may be noted that most amateur satellite communications are internationally coordinated and harmonized towards the lower end of the 2400-2450 MHz range in order to minimise ISM/WLAN interference.

For sharing studies 0 dBW Tx-power, 100 kHz bandwidth (BW), 0.5 dB feeder loss, and 3 dB antenna gain for a patch antenna are assumed for the satellite; and the amateur receiving ground station is assumed to be similar to the amateur service (100 kHz BW, 2dB noise figure). Table 61 summarizes the parameter choices for sharing studies.

The compatibility scenario considered in this section is the reception of a satellite downlink by an amateur radio ground station. MBANSs may be in the side-lobe of the amateur receiving station, for which 0dBi gain can be assumed.

Cumulative MBANS transmissions to an overhead LEO satellite uplink receiver are not studied, as it is reasonable to assume a proportionately low power density compared to existing WLANs/ISM equipment. Interference from Amateur TV into MBANSs is not studied either, since a preliminary assessment hinted at the absence of coexistence issues,

Table 61: Amateur Radio parameters

|  | Parameter | Value (Amateur Satellite) |
| --- | --- | --- |
| Amateur Transmitter | Bandwidth (MHz) | 0.100 |
|  | Max Tx power (dBm) | 30 |
|  | Antenna gain (dBi) | 0 |
|  | Antenna height (m) | n.a. |
| Amateur Receiver (Note 1) | Receiver sensitivity (dBm) | -112 |
|  | Antenna gain (dBi) | 20 (main beam), 0 (side-lobes) |
|  | Antenna height (m) | 10 |
|  | C/I objective (dB) | 20 |

Note 1: For the ground station in the downlink scenario

### Healthcare facility MBANSs

Table 52: MCL calculation in 2400 – 2440 MHz band between healthcare facility MBANSs and Amateur

| Victim characteristics | Units | Amateur Satellite |
| --- | --- | --- |
| Receiver bandwidth | MHz | 0.100 |
| Receiver noise figure | dB | 2 |
| Receiver antenna height | m | 10 |
| Receiver antenna gain | dBi | 0 |
| Operating frequency | MHz | 2420 |
| N, receiver thermal noise | dBm | -121.8 |
| I/N objective | dB | -10 |
| Interferer’s characteristics |  | MBANSs |
| e.i.r.p | dBm | 0 |
| Bandwidth | MHz | 3 |
| BW correction factor | dB | -14.8 |
| NFD (adjacent band interf) | dB | 0 |
| Wall attenuation | dB | 13 |
| Antenna height | m | 1.5 |
| Minimum path loss | dB | 104.1 |
| Interference distance FSL model | km | 1.57 |

Table 53: Interference from healthcare facility MBANSs to Amateur Radio - Settings and results

| Simulation input/output parameters | Settings / Results (Amateur Satellite) |
| --- | --- |
| Victim Link (VLK): Amateur |  |
| VLK frequency | 2420 MHz |
| VLK bandwidth | 100 kHz |
| VLR dRSS | -100 dBm |
| VLT → VLR path | FSL |
| Interfering Link (ILK): MBANS |  |
| ILK frequency | 2420 MHz |
| ILK bandwidth | 3 MHz |
| ILT Tx power | 0 dBm |
| ILT density | 40/km2 |
| ILT probability of transmission | 0.1 |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, above roof |
| ILT → VLR positioning mode | Closest interferer |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |
| iRSSunwanted, dBm (Std.dev., dB) | -157.98 (16.59) |
| Probability of interference (%) (C/I = 20 dB) | 2.6 |

### Home MBANSs

Table 55: MCL calculation in 2400 – 2440 MHz band between home MBANSs and Amateur

| Victim characteristics | Units | Amateur Satellite |
| --- | --- | --- |
| Receiver bandwidth | MHz | 0.100 |
| Receiver noise figure | dB | 2 |
| Receiver antenna height | m | 10 |
| Receiver antenna gain | dBi | 0 |
| Operating frequency | MHz | 2420 |
| N, receiver thermal noise | dBm | -121.8 |
| I/N objective | dB | -10 |
| Interferer’s characteristics |  | MBANSs |
| e.i.r.p | dBm | 13 |
| Bandwidth | MHz | 3 |
| BW correction factor | dB | -14.8 |
| NFD (adjacent band interf) | dB | 0 |
| Wall attenuation (Note 1) | dB | 0 (outdoor) |
| Antenna height | m | 1.5 |
| Minimum path loss | dB | 130.1 |
| Interference distance FSL model | km | 31.41 |

Note 1: Worst case configuration assumed, in which MBANS transmitter or receiver is outdoors.

Table 56: Interference from home MBANSs to Amateur Radio - Settings and results

| Simulation input/output parameters | Settings / Results (Amateur Satellite) |
| --- | --- |
| Victim Link (VLK): Amateur |  |
| VLK frequency | 2420 MHz |
| VLK bandwidth | 100 kHz |
| VLR dRSS | -100 dBm |
| VLT → VLR path | FSL |
| Interfering Link (ILK): MBANS |  |
| ILK frequency | 2420 MHz |
| ILK bandwidth | 3 MHz |
| ILT Tx power | 13 dBm |
| ILT density | 10/km2 |
| ILT probability of transmission | 0.02 |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, above roof |
| ILT → VLR positioning mode | Uniform density |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |
| iRSSunwanted, dBm (Std.dev., dB) | -158.67 (13.38) |
| Probability of interference (%) (C/I = 20 dB) | 1.0 |

### Ambulance MBANSs

Table 58: MCL calculation in 2400 – 2440 MHz band between ambulance MBANSs and Amateur

| Victim characteristics | Units | Amateur Satellite |
| --- | --- | --- |
| Receiver bandwidth | MHz | 0.100 |
| Receiver noise figure | dB | 2 |
| Receiver antenna height | m | 10 |
| Receiver antenna gain | dBi | 0 |
| Operating frequency | MHz | 2420 |
| N, receiver thermal noise | dBm | -121.8 |
| I/N objective | dB | -10 |
| Interferer’s characteristics |  | MBANSs |
| e.i.r.p | dBm | 0 |
| Bandwidth | MHz | 3 |
| BW correction factor | dB | -14.8 |
| NFD (adjacent band interf) | dB | 0 |
| Wall attenuation | dB | 7 |
| Antenna height | m | 1.5 |
| Minimum path loss | dB | 110.1 |
| Interference distance FSL model | km | 3.14 |

Table 59: Interference from ambulance MBANSs to Amateur Radio - Settings and results

| Simulation input/output parameters | Settings / Results (Amateur Satellite) |
| --- | --- |
| Victim Link (VLK): Amateur |  |
| VLK frequency | 2420 MHz |
| VLK bandwidth | 100 kHz |
| VLR dRSS | -100 dBm |
| VLT → VLR path | FSL |
| Interfering Link (ILK): MBANS |  |
| ILK frequency | 2420 MHz |
| ILK bandwidth | 3 MHz |
| ILT Tx power | 0 dBm |
| ILT density | 5/km2 |
| ILT probability of transmission | 0.1 |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, above roof |
| ILT → VLR positioning mode | Uniform density |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |
| iRSSunwanted, dBm (Std.dev., dB) | -155.96 (12.14) |
| Probability of interference (%) (C/I = 20 dB) | 1.3 |

### Summary Amateur and Amateur Satellite

[

The results of the study indicate co-channel coexistence between Amateur Satellite Systems in the 2400 – 2450 MHz range and MBANSs operating in healthcare facilities, homes, and ambulances. However, since terrestrial amateur services such as multimedia and Amateur TV are also present in the 2400 – 2450 MHz range, the results and conclusions of Section 4.2 are applicable to this section too. They can be found in Section 4.2.4.

]

## Adjacent broadband wireless systems (BWS)

As was mentioned in Section 4.3, LTE is expected to utilize the band 2300-2400 MHz, which is adjacent to the considered 2400-2483.5 MHz band. Since the involved power levels are quite high and there are currently no guard bands specified for the LTE, it is important to study also possible adjacent band interferences. This is done in Section 6.8, in which a very similar adjacent band coexistence case is presented. The same results and conclusion are applicable to this section.

# Compatibility Study for the Band 2483.5 – 2500 MHz

The compatibility studies presented in this section are structured in the following way:

- Section 6.x:  Introduction of the radio service / application analysed in that subsection, including all relevant technical parameters for the studies

- Section 6.x.1: Compatibility with healthcare facility MBANSs, including MCL calculations and SEAMCAT simulations

- Section 6.x.2: Compatibility with home MBANSs, including MCL calculations and SEAMCAT simulations

- Section 6.x.3: Compatibility with ambulance MBANSs mode, including MCL calculations and SEAMCAT simulations

- Section 6.x.4: Summary and conclusions for the analysed radio service / application

## Active Medical Implants

Low Power – Active Medical Implants (LP-AMI) are application-specific SRDs allowed to operate in the 2483.5 – 2500 MHz band. Due to the physical proximity of MBANS and LP-AMI devices (possibly on the same patient’s body) and their common usage scenarios (e.g. hospital wards, elderly care houses, and medical ambulatories), it was considered necessary to study the compatibility with LP-AMI.

LP-AMI is used for implantable device applications and related external telemetry medical products and is composed by two integral components: (1) LP-AMI implantable device and (2) Peripheral interrogator unit – LP-AMI-P. In general, LP-AMI may transmit only when queried by an LP-AMI-P device, a stationary device installed indoors. In practice, this restriction ensures indoor operation for all LP-AMI components.

The transmit power of battery-driven LP-AMI is anticipated to be lower than that of mains-driven LP-AMI-P. Additionally LP-AMI are attenuated some 20-30 dB or more due to body loss as reported in ETSI TR 102 655 [ ], when the implant is located 10 mm or deeper inside the patient’s body.

Based on ETSI TR 102 655 [ ] the LP-AMI was assumed to be implanted at 3 cm depth, with the corresponding body attenuation (27 dB). The wanted signal received from LP-AMI-P was taken as constant for a constant distance of 5 m between LP-AMI and LP-AMI-P. Some safety margin to tolerate LOS obstruction by people around the LP-AMI patient was also accounted for.

LP-AMI-P is in most cases required [ETSI EN 301 559-1] to use LBT and AFA when is transmitting. To model this functionality in SEAMCAT, a random uniform frequency channel selection within the 2484 MHz – 2499 MHz band was used. This approach was also utilized for the elaboration of ECC Report 149 [ ].

According to [ETSI EN 301 559-1], when a Medical Implant Event occurs, the LP-AMI implant may immediately transmit time critical data associated with that Medical Implant Event to a LP-AMI-P without regard to channel occupancy. This operation mode—without utilizing LBT and AFA—is limited to 0.83% of any 1-hour period. Moreover it is limited to emergency situations, such as heart attacks, Medical Implant Events are treated differently in subsections 6.1.1, 6.1.2, and 6.1.3, due to the differences in the operational environment.

The main system parameters for LP-AMI are presented in Table 65.

Table 61: LP-AMI parameters

|  | Parameter | Value |
| --- | --- | --- |
| LP-AMI transmitter | Bandwidth (MHz) | Usually1. Occasionally 3. |
|  | Frequency distribution (MHz) | Uniform (2484-2499) with step 1 |
|  | Max Tx power (dBm) | 10 |
|  | Antenna gain (dBi) | 0 |
|  | Antenna height (m) | 1.5 |
| LP-AMI receiver | User defined dRSS (dBm/MHz) | -80 |
|  | Antenna gain (dBi) | 0 |
|  | Antenna height (m) | 1.5 |
|  | C/I objective (dB) | 12 |

### Healthcare facility MBANSs

Table 66 shows the MCL calculations for the minimum interference distances between LP-AMI-P and MBANS (both for same and next room) and between LP-AMI and MBANS (same room). Since the obtained distances do not guarantee coexistence, a more detailed analysis via SEAMCAT simulations is presented in the remainder of the subsection. LP-AMI may use the aggregation of 3 channels (equivalent to 1 channel of 3 MHz) for download sessions of recorded data. This mode of communication should not influence the DC of LP-AMI and is referred to as ‘turbo mode’ in the remainder of the report. The turbo mode is expected to be used mainly in hospitals. Hence this additional case is considered in the simulations of this subsection.

Table 62: MCL calculation in 2483,5 – 2500 MHz band between healthcare facility MBANSs and 
LP-AMI

| Victim characteristics | Units | LP-AMI-P |  | MBANSs |  | LP-AMI | MBANSs |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1 |  | 3 |  | 1 | 3 |
| Receiver noise figure | dB | 10 |  | 10 |  | 10 | 10 |
| Receiver antenna height | m | 1.5 |  | 1.5 |  | 1.5 | 1.5 |
| Receiver antenna gain | dBi | 0 |  | 0 |  | 0 | 0 |
| Operating frequency | MHz | 2490 |  | 2490 |  | 2490 | 2490 |
| N, receiver thermal noise | dBm | -103.8 |  | -99.1 |  | -103.8 | -99.1 |
| I/N objective | dB | 0 |  | 0 |  | 0 | 0 |
| Interferer’s characteristics |  | MBANSs |  | LP-AMI-P |  | MBANSs | LP-AMI |
| e.i.r.p | dBm | 0 |  | 10 |  | 0 | 10 |
| Bandwidth | MHz | 3 |  | 1 |  | 3 | 1 |
| BW correction factor | dB | -4.77 |  | 0 |  | -4.77 | 0 |
| NFD (adjacent band interf) | dB | 0 |  | 0 |  | 0 | 0 |
| Wall attenuation | dB | 10(next room) | 0 | 10(next room) | 0 | 27(body) | 27(body) |
| Antenna height | m | 1.5 |  | 1.5 |  | 1.5 | 1.5 |
| Minimum path loss | dB | 89.1 | 99.1 | 99.1 | 109.1 | 72.1 | 82.1 |
| Interference distance FSL model | km | 0.27 | 0.86 | 0.86 | 2.72 | 0.04 | 0.12 |

#### Interference from MBANSs to LP-AMI-P

The considered scenarios (next and same room) are depicted in Figure 38 and Figure 39 correspondingly. For the neighbouring room case, an attenuation of 10 dB is applied to the free space propagation model as result of the indoor wall. The simulations analyse several point-to-point interference scenarios at different MBANS-LP-AMI separation distances. Medical Implant Events are not considered in this healthcare facility scenario. The settings for all cases and their results are summarized in Table 67.

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

10 dB building penetration loss

Healthcare facility

Wanted signal paths

Interference paths

LP–AMI-P

LP-AMI

5 m

Figure 41: Interference scenario – healthcare facility MBANS into LP-AMI-P (next room)

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

Healthcare facility

Wanted signal paths

Interference paths

LP–AMI-P

LP-AMI

5 m

Figure 42: Interference scenario – healthcare facility MBANS into LP-AMI-P (same room)

Table 63: Interference from healthcare facility MBANS to LP-AMI-P - Settings and results

| Simulation input/output parameters | Settings / Results |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Victim Link (VLK): LP-AMI TO LP-AMI-P |  |  |  |  |  |  |
| VLK frequency | Uniform 2484-2499 MHz |  |  |  |  |  |
| VLK bandwidth | 1 MHz (normal mode) |  |  | 3 MHz (turbo mode) |  |  |
| VLR dRSS | -80 dBm/1 MHz |  |  | -75.3 dBm/3 MHz |  |  |
| Interfering Link (ILK): MBANS |  |  |  |  |  |  |
| ILK frequency | 2490 MHz |  |  |  |  |  |
| ILK bandwidth | 3 MHz |  |  |  |  |  |
| ILT Tx power | 0 dBm (10% probability) |  |  |  |  |  |
| ILT → VLR positioning mode | Correlated |  |  |  |  |  |
| ILT → VLR interfering path | FSL |  |  |  |  |  |
| ILT → VLR distance | 1 m | 5 m | 10 m | 1 m | 5 m | 10 m |
| Probability of interference (%) (same room) (C/I = 12 dB) | 10.2 | 10.1 | - | 10.1 | 9.9 | - |
| Probability of interference (%) (next room) (C/I = 12 dB) | - | 7.8 | 2.7 | - | 8.3 | 4.6 |

#### Interference from MBANSs to LP-AMI

. [the effect on LP-AMI battery life is another matter, currently under consideration by the LP-AMI industry].A worst-case scenario was considered in which LP-AMI and MBANSs are used simultaneously on the same user (patient) and in the same part of the body. In such scenario—depicted in Figure 40—the MBANS transmitter is located directly over the LP-AMI device, which is implanted at 3 cm depth. An attenuation of 27.04 dB was applied to the free space propagation model to model the attenuation caused by the body tissue. The SEAMCAT parameters of the MBANS wearable transmitter and the LP-AMI receiver are identical to those shown in Table 67 except for the interfering path, which is FSL with an offset -27.04 dB, and the ILT-VLR distance, which is 3 cm. The resulting interference probability is 10%, both for normal and turbo mode.

MBANS

MBANS

hub

e.i.r.p= 0 dBm

Healthcare facility

Wanted signal paths

Interference paths

LP-AMI-P

LP-AMI

5 m

Figure 43: Interference scenario – healthcare facility MBANS into LP-AMI

#### Interference from LP-AMI-P to MBANSs

The considered interference scenarios are depicted in Figure 41 and Figure 42. The simulation approach is equivalent to that used in Section 6.1.1.1. The simulation settings and results are summarized in Table 68.

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

10 dB building penetration loss

Healthcare facility

Wanted signal paths

Interference paths

LP–AMI-P

LP-AMI

Figure 44: Interference scenario – LP-AMI-P into healthcare facility MBANS (next room)

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

Healthcare facility

Wanted signal paths

Interference paths

LP–AMI-P

LP-AMI

Figure 45: Interference scenario – LP-AMI-P into healthcare facility MBANS (same room)

Table 64 : Interference from LP-AMI-P to healthcare facility MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |  |  |  |  |
| VLK frequency | 2490 MHz |  |  |  |  |  |
| VLK bandwidth | 3 MHz |  |  |  |  |  |
| VLT → VLR path | FSL (user-defined radius, 3 m) |  |  |  |  |  |
| Interfering Link (ILK): LP-AMI-P TO LP-AMI |  |  |  |  |  |  |
| ILK frequency | Uniform 2484-2499 MHz |  |  |  |  |  |
| ILK bandwidth | 1 MHz (normal mode) |  |  | 3 MHz (turbo mode) |  |  |
| ILT Tx power | 10 dBm (10% probability) |  |  |  |  |  |
| ILT → VLR positioning mode | Correlated |  |  |  |  |  |
| ILT → VLR interfering path | FSL |  |  |  |  |  |
| ILT → VLR distance | 1 m | 5 m | 10 m | 1 m | 5 m | 10 m |
| Probability of interference (%) (same room) (C/I = 15 dB) | 3 | 2.7 | - | 4.9 | 4.8 | - |
| Probability of interference (%) (next room) (C/I = 15 dB) | - | 2.2 | 1.3 | - | 4 | 2.8 |

#### Interference from LP-AMI to MBANSs

Similarly to Section 6.1.1.2, a worst-case scenario is considered in which LP-AMI and MBANSs are used simultaneously on the same user (patient) and in the same part of the body. The scenario is depicted in Figure 43.

The parameters of the MBANS wearable transmitter and the LP-AMI receiver are identical to those in Table 68, except for the interfering path, which is FSL with an offset -27.04 dB, and the MBANS-LP-AMI separation distance, which is 3 cm. The resulting interference probability is 3.4%. For the turbo mode, when LP-AMI uses 3 MHz bandwidth, the interference probability is 5.7 %.

MBANS

MBANS

hub

e.i.r.p= 0 dBm

Healthcare facility

Wanted signal paths

Interference paths

LP-AMI-P

LP-AMI

Figure 46: Interference scenario – LP-AMI into healthcare facility MBANS

### Home MBANSs

Table 68 shows the MCL calculations for the minimum interference distances between LP-AMI-P and home MBANS (both for same and next rooms) and between LP-AMI and home MBANS (same room). Only the one direction of interference is presented; MBANS on LP-AMI(-P). The complementary calculations—LP-AMI(-P) on MBANSs—are the same as for the healthcare facility case presented in Table 62, since the MBANS interference criterion remains unaffected. The obtained distances do not guarantee coexistence. Therefore a more detailed analysis via SEAMCAT simulations is presented in the remainder of the subsection.

Table 68: MCL calculation in 2483.5 – 2500 MHz band between home MBANSs and LP-AMI (P)

| Victim characteristics | Units | LP-AMI-P |  | LP-AMI |
| --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1 |  | 1 |
| Receiver noise figure | dB | 10 |  | 10 |
| Receiver antenna height | m | 1.5 |  | 1.5 |
| Receiver antenna gain | dBi | 0 |  | 0 |
| Operating frequency | MHz | 2490 |  | 2490 |
| N, receiver thermal noise | dBm | -103.8 |  | -103.8 |
| I/N objective | dB | 0 |  | 0 |
| Interferer’s characteristics |  | MBANSs |  | MBANSs |
| e.i.r.p | dBm | 13 |  | 13 |
| Bandwidth | MHz | 3 |  | 3 |
| BW correction factor | dB | -4.77 |  | -4.77 |
| NFD (adjacent band interf) | dB | 0 |  | 0 |
| Wall attenuation | dB | 10(next room) | 0 | 27(body) |
| Antenna height | m | 1.5 |  | 1.5 |
| Minimum path loss | dB | 102.1 | 112.1 | 85.1 |
| Interference distance FSL model | km | 1.22 | 3.84 | 0.17 |

#### Interference from MBANSs to LP-AMI-P

The assumed scenarios are the same as in Section 6.1.1.1.. In this case the turbo mode is not considered in the analysis. The simulation parameters and the results for the point-to-point interference scenarios are summarized in Table 70.

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

10 dB building penetration loss

Home

Wanted signal paths

Interference paths

LP–AMI-P

LP-AMI

5 m

Figure 47: Interference scenario – home MBANS into LP-AMI-P (next room)

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Home

Wanted signal paths

Interference paths

LP–AMI-P

LP-AMI

5 m

Figure 48: Interference scenario – home MBANS into LP-AMI-P (same room)

Table 66: Interference from home MBANS to LP-AMI-P - Settings and results

| Simulation input/output parameters | Settings / Results |  |  |
| --- | --- | --- | --- |
| Victim Link (VLK): LP-AMI TO LP-AMI-P |  |  |  |
| VLK frequency | Uniform 2484-2499 MHz |  |  |
| VLK bandwidth | 1 MHz |  |  |
| VLR dRSS | -80 dBm/1 MHz |  |  |
| Interfering Link (ILK): MBANS |  |  |  |
| ILK frequency | 2490 MHz |  |  |
| ILK bandwidth | 3 MHz |  |  |
| ILT Tx power | 13 dBm |  |  |
| ILT → VLR positioning mode | Correlated |  |  |
| ILT → VLR interfering path | FSL |  |  |
| ILT → VLR distance | 1 m | 5 m | 10 m |
| Probability of interference (%) (same room) (C/I = 12 dB) | 2.1 | 2 | - |
| Probability of interference (%) (next room) (C/I = 12 dB) | - | 1.7 | 1.9 |

Note

1: Medical Implants Events have also been considered by changing the victim link (VLK) central frequency to a fixed value of 2490 MHz. The obtained interference probability values do not differ and are hence not explicitly presented.

#### Interference from MBANSs to LP-AMI

Here the same justification applies as in case 6.1.1.2. The resulting interference probability is 2.1 %,

MBANS

MBANS

hub

e.i.r.p= 13 dBm

Home

Wanted signal paths

Interference paths

LP-AMI-P

LP-AMI

5 m

Figure 49: Interference scenario – home MBANS into LP-AMI

#### Interference from LP-AMI-P to MBANSs

The assumed scenarios are the complementary to those of Section 6.1.2.1 and are depicted in Figure 47 and Figure 48. The simulation parameters and the corresponding results are in Table 71.

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

10 dB building penetration loss

Location independent

Wanted signal paths

Interference paths

LP–AMI-P

LP-AMI

Figure 50: Interference scenario – LP-AMI-P into home MBANS (next room)

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Location independent

Wanted signal paths

Interference paths

LP–AMI-P

LP-AMI

Figure 51: Interference scenario – LP-AMI-P into home MBANS (same room)

Table 67: Interference from LP-AMI-P to home MBANSs - Settings and results

| Simulation input/output parameters | Settings / Results |  |  |
| --- | --- | --- | --- |
| Victim Link (VLK): MBANS |  |  |  |
| VLK frequency | 2490 MHz |  |  |
| VLK bandwidth | 3 MHz |  |  |
| VLT → VLR path | FSL (user-defined radius, 10 m) |  |  |
| Interfering Link (ILK): LP-AMI-P TO LP-AMI |  |  |  |
| ILK frequency | Uniform 2484-2499 MHz |  |  |
| ILK bandwidth | 1 MHz |  |  |
| ILT Tx power | 10 dBm |  |  |
| ILT → VLR positioning mode | Correlated |  |  |
| ILT → VLR interfering path | FSL |  |  |
| ILT → VLR distance | 1 m | 5 m | 10 m |
| Probability of interference (%) (same room) (C/I = 15 dB) | 2.8 | 2.6 | - |
| Probability of interference (%) (next room) (C/I = 15 dB) | - | 2 | 0.5 |

#### Interference from LP-AMI to MBANSs

Here the same justification applies as in Section 6.1.1.4. The resulting interference is 3.7 %.

MBANS

MBANS

hub

e.i.r.p= 13 dBm

Home

Wanted signal paths

Interference paths

LP-AMI-P

LP-AMI

Figure 52: Interference scenario – LP-AMI into home MBANS

### Ambulance MBANSs

Active medical implant communication is only specified inside buildings. Therefore it is considered unnecessary to analyse compatibility with ambulance MBANSs in further detail. The physical distance between in-building active medical implants and in-ambulance MBANSs—together with the attenuation added by the building wall and ambulance chassis—provide enough evidence to expect good coexistence.

### Summary Active Medical Implants

[

[In this subsection the text should also mention that future ETSI standards of LP-AMI and MBANSs might specify some means of detecting each other for improving coexistence, especially with regard to LP-AMI implants‘ battery life.] [A liason statement was agreed to be sent on Jan 8 by SE24 to ETSI TG30, requesting for their input on this issue and the current results and conclusions of section 6.1, presented directly below]

Healthcare facility MBANSs:

According to the results presented above, in worst-case situations—modelled with MCL calculations—the interference distance between active medical implants (LP-AMI) and MBANSs is 40 m - 860 m (MBANS as interferer) and 120 m – 2.7 km (LP-AMI as interferer). The simulation-based analysis yielded 2.7% - 8.3% probability of interference from MBANSs into LP-AMI, when both systems are located in a neighbouring room and up to a distance of 10 m from one another. When operated in the same room or on the same patient, the probability of interference from MBANSs into LP-AMI raises to approximately 10%. In the other coexistence direction, the probability of interference from LP-AMI into MBANSs lies in the ranges of 1.3% - 4% (neighbouring room and up to a distance of 10 m), 2.7% - 4.9% (same room), and 3.4% - 5.7% (same patient).

Co-channel compatibility between MBANSs and LP-AMI in healthcare facilities is not given, since significant interference levels are expected, even when both systems are operated in different rooms.

Home MBANSs:

In worst-case situations—modelled with MCL calculations—the interference distance between active medical implants (LP-AMI) and MBANSs is 170 m – 3.8 km (MBANS as interferer). The simulation-based analysis yielded around 2% probability of interference from MBANSs into LP-AMI, when both systems are located in the same room or on the same patient. In the other coexistence direction, the probability of interference from LP-AMI into MBANSs lies in the range of 2.6% - 3.7%, when both systems are located in the same room or on the same patient.

Co-channel compatibility between MBANSs and LP-AMI in patient homes appears possible, based on the obtained simulation results.

Ambulance MBANS:

Ambulance MBANSs were not explicitly considered, since LP-AMI communication is not expected within ambulance. Co-channel compatibility between MBANSs and LP-AMI in patient homes is expected.]

## ISM

ISM devices are left out of the study, as explained in Section 5.2.

## MSS Systems (Globalstar mobile phones)

The band 2483.5 – 2500 MHz is used for MSS communications in the direction Space-to-Earth, paired with 1610 – 1626.5 MHz for transmissions Earth-to-Space. Today this MSS allocation is being utilised by the Globalstar system, which is based on CDMA IS-95 technology and serves around 300000 reported subscribers worldwide.

The danger of interference from MSS into MBANSs may be discarded based on the following analysis:

- Globalstar downlink signal should comply with maximum Pfd limit (for single satellite) of -124.5 dBW/m2/MHz on the ground,

- Effective antenna area at 2483.5 MHz is -30.2 dB(m2), MBANS receiver bandwidth 5 MHz,

- Resulting MSS downlink signal power in MBANS receiver is -147.7 dBW= -117.7 dBm,

- Even considering power summation from multiple Globalstar satellites, the resulting interfering power compares favourably with the noise floor of MBANS receiver of -91 dBm.

It thus may be concluded that MSS downlink should not pose any danger to MBANS operations. Therefore it was decided to limit the statistical study only to the case of interference from MBANSs to Globalstar MES (Mobile Earth Station) receivers, as considered by previous studies [ECC Report 149]. The relevant parameters for Globalstar MES to be used in statistical study were taken from the ECC Report 165 [ ].

### Healthcare facility MBANSs

Table 68: MCL calculation in 2483.5 – 2500 MHz band between healthcare facility MBANSs and MSS

| Victim characteristics | Units | MSS MES |  |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1.23 |  |
| Receiver noise figure | dB | 4 |  |
| Receiver antenna height | m | 1.5 |  |
| Receiver antenna gain | dBi | 0 |  |
| Operating frequency | MHz | 2490 |  |
| N, receiver thermal noise | dBm | -108.93 |  |
| I/N objective | dB | -12 |  |
| Interferer’s characteristics |  | MBANSs |  |
| e.i.r.p | dBm | 0 (single MBANS) | 13 (additive contribution of 20 MBANSs) |
| Bandwidth | MHz | 3 |  |
| BW correction factor | dB | -3.87 |  |
| NFD (adjacent band interf) | dB | 0 |  |
| Wall attenuation | dB | 13 |  |
| Antenna height | m | 1.5 |  |
| Minimum path loss | dB | 104.1 | 117.1 |
| Interference distance FSL model | km | 1.53 | 6.84 |
| Interference distance P.452 model with 50% probability | km | 0.61 | 1.53 |

Table 68 presents interference distances considering the impact of a single MBANSs and the impact of several MBANSs (area with high density). The latter is calculated assuming that a significant part of the MBANSs concentrated in a hospital are next to the wall that separates the hospital from the street on which the MSS MES (MSS terminal) is located. The assumed 20 MBANSs are not shielded by the hospital inner clutter and their e.i.r.p values are added.

Using the propagation model described in ITU-R recommendation P.452 []—which is valid for point-point propagation on the Earth surface for frequencies above 0.7 GHz—the minimum distances are lower than using FSL, as shown in the last row of Table 68. The propagation distance was calculated with the P.452 model with a probability of 50%. A probability of 50% means that the propagation loss could be lower than the given value 50% of the time and should be higher or equal 50% of the time. Therefore the given distance will be sufficient or more than needed at least 50% of the time.

#### Interference from MBANSs to Globalstar MES receivers

The considered scenario is depicted in Figure 50 and simulation settings are in Table 73.

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Victim MES receiver

Wanted signal paths

Interference paths

Figure 53: Interference scenario – healthcare facility MBANS into MSS MES

Table 69: Interference from healthcare facility MBANS to MSS MES - Settings and results

| Simulation input/output parameters | Settings / Results |
| --- | --- |
| Victim Link (VLK): MSS SATELLITE TO MSS MES |  |
| VLK frequency | 2490 MHz |
| VLK bandwidth | 1.23 MHz |
| VLR dRSS | -100 dBm |
| VLR noise floor | -109 dBm |
| Interfering Link (ILK): MBANS |  |
| ILK frequency | 2490 MHz |
| ILK bandwidth | 3 MHz |
| ILT Tx power | 0 dBm |
| ILT density | 40/km2 |
| ILT probability of transmission | 0.1 |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, below roof |
| ILT → VLR positioning mode | Closest interferer |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |
| iRSSunwanted, dBm (Std.dev., dB) | -166.19 (19.7) |
| Probability of interference (%) (I/N = -12 dB) | 3.5 |

### Home MBANSs

Table 70: MCL calculation in 2483.5 – 2500 MHz band between home MBANSs and MSS

| Victim characteristics | Units | MSS MES |
| --- | --- | --- |
| Receiver bandwidth | MHz | 1.23 |
| Receiver noise figure | dB | 4 |
| Receiver antenna height | m | 1.5 |
| Receiver antenna gain | dBi | 0 |
| Operating frequency | MHz | 2490 |
| N, receiver thermal noise | dBm | -108.93 |
| I/N objective | dB | -12 |
| Interferer’s characteristics |  | MBANSs |
| e.i.r.p | dBm | 13 |
| Bandwidth | MHz | 3 |
| BW correction factor | dB | -3.87 |
| NFD (adjacent band interf) | dB | 0 |
| Wall attenuation | dB | 0 (outdoor) (Note 1) |
| Antenna height | m | 1.5 |
| Minimum path loss | dB | 130.1 |
| Interference distance FSL model | km | 30.53 |
| Interference distance P.452 model with 50% probability | km | 4.38 |

Note 1: Worst case configuration assumed, in which MBANS transmitter or receiver is outdoors.

According to the MCL calculations, home MBANSs could interfere with MSS receivers located less than 4.8 km around the MBANS. Since such a distance was regarded problematic the probability of interference is analysed in further detail in the following subsections.

#### Interference from MBANSs to Globalstar MES receivers

Scenario is depicted in Figure 52 and simulation settings are in Table 75

Victim MES receiver

Wanted signal paths

Interference paths

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Patient home

13 dB building penetration loss

Figure 55: Interference scenario – home MBANS into MSS MES

Table 71: Interference from home MBANS to MSS MES - Settings and results

| Simulation input/output parameters | Settings / Results |
| --- | --- |
| Victim Link (VLK): MSS SATELLITE TO MSS MES |  |
| VLK frequency | 2490 MHz |
| VLK bandwidth | 1.23 MHz |
| VLR dRSS | -100 dBm |
| VLR noise floor | -109 dBm |
| Interfering Link (ILK): MBANS |  |
| ILK frequency | 2490 MHz |
| ILK bandwidth | 3 MHz |
| ILT Tx power | 13 dBm (2% probability) |
| ILT density | 10/km2 |
| ILT probability of transmission | 1 |
| ILT → VLR interfering path | Extended Hata, urban, indoor→outdoor, below roof |
| ILT → VLR positioning mode | Uniform density |
| ILT → VLR number of transmitters | 785 (for 5 km simulation radius) |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |
| iRSSunwanted, dBm (Std.dev., dB) | -158.68 (10.43) |
| Probability of interference (%) (I/N = -12 dB) | 0.9 |

### Ambulance MBANSs

Table 72: MCL calculation in 2483,5 – 2500 MHz band between healthcare facility MBANSs and MSS

| Victim characteristics | Units | MSS MES |
| --- | --- | --- |
| Receiver bandwidth | MHz | 1.23 |
| Receiver noise figure | dB | 4 |
| Receiver antenna height | m | 1.5 |
| Receiver antenna gain | dBi | 0 |
| Operating frequency | MHz | 2490 |
| N, receiver thermal noise | dBm | -108.93 |
| I/N objective | dB | -12 |
| Interferer’s characteristics |  | MBANSs |
| e.i.r.p | dBm | 0 |
| Bandwidth | MHz | 3 |
| BW correction factor | dB | -3.87 |
| NFD (adjacent band interf) | dB | 0 |
| Wall attenuation | dB | 7 (ambulance) |
| Antenna height | m | 1.5 |
| Minimum path loss | dB | 110.1 |
| Interference distance FSL model | km | 3.05 |
| Interference distance P.452 model with 50% probability | km | 0.91 |

#### Interference from MBANSs to Globalstar MES receivers

Scenario is depicted in Figure 53 and simulation settings are on Table 77.

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB   penetration loss

Wanted signal paths

Interference paths

Victim MES receiver

Figure 56: Interference scenario – ambulance MBANS into MSS MES

Table 73: Interference from ambulance MBANS to MSS MES- Settings and results

| Simulation input/output parameters | Settings / Results |
| --- | --- |
| Victim Link (VLK): MSS SATELLITE TO MSS MES |  |
| VLK frequency | 2490 MHz |
| VLK bandwidth | 1.23 MHz |
| VLR dRSS | -100 dBm |
| VLR noise floor | -109 dBm |
| Interfering Link (ILK): MBANS |  |
| ILK frequency | 2490 MHz |
| ILK bandwidth | 3 MHz |
| ILT Tx power | 0 dBm |
| ILT density | 5/km2 |
| ILT probability of transmission | 0.1 |
| ILT → VLR interfering path | Extended Hata, urban, intdoor→outdoor, below roof |
| ILT → VLR positioning mode | Uniform density |
| ILT → VLR number of active transmitters | 2 (for 1 km simulation radius) |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |
| iRSSunwanted, dBm (Std.dev., dB) | -169.1 (14.6) |
| Probability of interference (%) (I/N = -12 dB) | 1.4 |

### Summary MSS Systems

[

Healthcare facility MBANSs:

According to the results presented above, in worst-case situations—modelled with MCL calculations—the interference distance between MSS and MBANSs is 0.6 km - 1.5 km (single MBANS as interferer) and 1.5 km - 6.8 km (multiple MBANSs as interferers). The simulation-based average-case analysis yielded 3.5% interference probability from MBANSs into MSS.

MSS systems operating in the 2483.5 — 2500 MHz range and healthcare facility MBANSs operating in the same range appear to be compatible.

Home and ambulance MBANSs:

In worst-case situations—modelled with MCL calculations—the interference distance between MSS and MBANSs is 4.4 km - 31 km (home MBANS as interferer) and 900 m - 3 km (ambulance MBANS as interferer). The simulation-based average-case analysis yielded an interference probability into MSS of 0.9% (from home MBANSs) and 1.4% (from ambulance MBANSs).

MSS systems operating in the 2483.5 — 2500 MHz range and home and ambulance MBANSs operating in the same range appear to be compatible.

]

## CGC (Complementary Ground Component of MSS systems)

The CGC is an emerging future possible idea of supplementary evolution of the MSS networks whereas terrestrial base stations would be installed in order to improve the coverage of MSS signals, e.g. within conditions of dense urban environments where very low sky observation angles severely hamper reliable reception of MSS satellite signals (so called “city canyon” scenario). A detailed description of intended operation of CGC as a part of Globalstar system may be found in ECC Report 165 [ ].

For systems to be considered in the studies related to the introduction of CGC associated with non GSO MSS systems in the bands 1.6 and 2.5 GHz, it was decided to limit the studies to the GLOBALSTAR case in the band 2483.5 – 2500 MHz and IRIDIUM and GLOBALSTAR in the band 1610 – 1626.5 MHz, for which parameters were received by CEPT.

An important feature of CGC to be considered in this study is that CGC would be operated in a portion of the same frequency bands as their satellite-based mother-systems. In other words, the CGC for Globalstar system deployed in the frequency band 2483.5 – 2500 MHz would also operate in the same band but just on a sub-set of available radio channels. Therefore when analysing CGC as part of co-existence analysis with MBANS applications, we still need to consider the same MSS MES device as victim receiver, but now we need to consider CGC BS (Base Station) emissions as potential interferer to the MBANS receiver. Previous studies [ECC Report 149] have selected CGC BS transmitters as the largest source of interference. The required parameters of CGC BS emissions are taken from ECC Report 165 [ ].  It should be also noted that, when used with CGC, the MES will have better link budget and could be therefore operated indoors.

### Healthcare facility MBANSs

Table 74: MCL calculation in 2483.5 – 2500 MHz band between healthcare facility MBANSs and MSS-CGC

| Victim characteristics | Units | CGC MES | MBANSs |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1.23 | 3 |
| Receiver noise figure | dB | 4 | 10 |
| Receiver antenna height | m | 1.5 | 1.5 |
| Receiver antenna gain | dBi | 0 | 0 |
| Operating frequency | MHz | 2490 | 2490 |
| N, receiver thermal noise | dBm | -108.93 | -99.1 |
| I/N objective | dB | -12 | 0 |
| Interferer’s characteristics |  | MBANSs | CGC BS Tx |
| e.i.r.p | dBm | 0 | 17 |
| Bandwidth | MHz | 3 | 1.23 |
| BW correction factor | dB | -3.87 | 0 |
| NFD (adjacent band interf) | dB | 0 | 0 |
| Wall attenuation | dB | n.a. (indoor) | 13 |
| Antenna height | m | 1.5 | 30 |
| Minimum path loss | dB | 117.1 | 103.1 |
| Interference distance FSL model | km | 6.84 | 1.36 |
| Interference distance P.452 model with 50% probability | km | 1.53 | 1.36 |

#### Interference from MBANSs to Globalstar MES receivers in CGC mode

Wanted signal paths

Interference paths

13 dB building penetration loss

e.i.r.p= 0 dBm

Healthcare facility

CGC BS

MBANS

sensor

Outdoor victim MES receiver

MBANS

hub

Indoor victim MES receiver

Figure 57: Interference scenario – healthcare facility MBANS into CGC MES

Table 75: Interference from healthcare facility MBANS to CGC MES - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): CGC BASE STATION TO CGC MES |  |  |
| VLK frequency | 2490 MHz |  |
| VLK bandwidth | 1.23 MHz |  |
| VLR dRSS | -100 dBm |  |
| VLR noise floor | -109 dBm |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2490 MHz |  |
| ILK bandwidth | 3 MHz |  |
| ILT Tx power | 0 dBm |  |
| ILT density | 40/km2 |  |
| ILT probability of transmission | 0.1 |  |
| ILT → VLR interfering path | Extended Hata, urban, below roof |  |
|  | indoor→indoor | indoor→outdoor |
| ILT → VLR positioning mode | Closest interferer |  |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) | -100 (0) |
| iRSSunwanted, dBm (Std.dev., dB) | -178.3 (21. 42) | -166.19 (19.7) |
| Probability of interference (%) (I/N = -12 dB) | 1.8 | 3.5 |

#### Interference from CGC base station to MBANSs

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Outoor wanted MES receiver

Wanted signal paths

Interference paths

Indoor wanted MES receiver

CGC BS

Figure 58: Interference scenario – GC BS into healthcare facility MBANS

Table 76: Interference from CGC BS to healthcare facility MBANS - Settings and results

| Simulation input/output parameters | Settings / Results |
| --- | --- |
| Victim Link (VLK): MBANS |  |
| VLK frequency | 2490 MHz |
| VLK bandwidth | 3 MHz |
| VLT → VLR path | FSL (user-defined radius, 3 m) |
| Interfering Link (ILK): CGC link (BS to MES) |  |
| ILK frequency | 2490 MHz |
| ILK bandwidth | 5 MHz |
| ILT Tx power | 43 dBm |
| ILT peak antenna gain | 19 dBi |
| ILT density | 0.02/km2 |
| ILT probability of transmission | 1 |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof |
| ILT → VLR positioning mode | Uniform density |
| ILT → VLR number of active transmitters | 1 |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -45.61 (4.39) |
| iRSSunwanted, dBm (Std.dev., dB) | -103.23 (12.9) |
| Probability of interference (%) (C/I = 15 dB) | 0.4 |

### Home MBANSs

Table 77: MCL calculation in 2483.5 – 2500 MHz band between home MBANSs and MSS-CGC

| Victim characteristics | Units | CGC MES | MBANSs |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1.23 | 3 |
| Receiver noise figure | dB | 4 | 10 |
| Receiver antenna height | m | 1.5 | 1.5 |
| Receiver antenna gain | dBi | 0 | 0 |
| Operating frequency | MHz | 2490 | 2490 |
| N, receiver thermal noise | dBm | -108.93 | -99.1 |
| I/N objective | dB | -12 | 0 |
| Interferer’s characteristics |  | MBANSs | CGC BS Tx |
| e.i.r.p | dBm | 13 | 17 |
| Bandwidth | MHz | 3 | 1.23 |
| BW correction factor | dB | -3.87 | 0 |
| NFD (adjacent band interf) | dB | 0 | 0 |
| Wall attenuation (Note 1) | dB | 0 (indoor) | 0 (outdoor) |
| Antenna height | m | 1.5 | 30 |
| Minimum path loss | dB | 130.1 | 116.1 |
| Interference distance FSL model | km | 30.53 | 6.09 |
| Interference distance P.452 model with 50% probability | km | 4.38 | 5.24 |

Note1: Worst case configuration assumed, in which MBANS transmitter or receiver is outdoors.

#### Interference from MBANSs to Globalstar MES receivers in CGC mode

Wanted signal paths

Interference paths

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

Home

13 dB building penetration loss

Indoor victim MES receiver

CGC BS

Figure 59: Interference scenario – home MBANS into CGC MES

Table 78: Interference from home MBANS to CGC MES - Settings and results

| Simulation input/output parameters | Settings / Results |  |
| --- | --- | --- |
| Victim Link (VLK): CGC BASE STATION TO CGC MES |  |  |
| VLK frequency | 2490 MHz |  |
| VLK bandwidth | 1.23 MHz |  |
| VLR dRSS | -100 dBm |  |
| VLR noise floor | -109 dBm |  |
| Interfering Link (ILK): MBANS |  |  |
| ILK frequency | 2490 MHz |  |
| ILK bandwidth | 3 MHz (2% probability) |  |
| ILT Tx power | 13 dBm |  |
| ILT density | 10/km2 |  |
| ILT probability of transmission | 1 |  |
| ILT → VLR interfering path | Extended Hata, urban, below roof |  |
|  | indoor→indoor | indoor→outdoor |
| ILT → VLR positioning mode | Uniform |  |
| ILT → VLR number of transmitters | 785 (for 5 km simulation radius) |  |
| Simulation results |  |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) | -100 (0) |
| iRSSunwanted, dBm (Std.dev., dB) | -178.28 (16.44) | 158.68 (10.43) |
| Probability of interference (%) (I/N = -12 dB) | 0.6 | 0.9 |

#### Interference from CGC base station to MBANSs

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Outoor wanted MES receiver

Wanted signal paths

Interference paths

Indoor wanted MES receiver

CGC BS

Figure 60: Interference scenario – CGC BS into home MBAN

Table 79: Interference from CGC BS to home MBANS - Settings and results

| Simulation input/output parameters | Settings / Results |
| --- | --- |
| Victim Link (VLK): MBANS |  |
| VLK frequency | 2490 MHz |
| VLK bandwidth | 3 MHz |
| VLT → VLR path | FSL (user-defined radius, 10 m) |
| Interfering Link (ILK): CGC link (BS to MES) |  |
| ILK frequency | 2490 MHz |
| ILK bandwidth | 5 MHz |
| ILT Tx power | 43 dBm |
| ILT peak antenna gain | 19 dBi |
| ILT density | 0.02/km2 |
| ILT probability of transmission | 1 |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof |
| ILT → VLR positioning mode | Uniform density |
| ILT → VLR number of active transmitters | 1 |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -44.2 (4.44) |
| iRSSunwanted, dBm (Std.dev., dB) | -103.84 (12.92) |
| Probability of interference (%) (C/I = 15 dB) | 0.3 |

### Ambulance MBANSs

Table 80: MCL calculation in 2483.5 – 2500 MHz band between ambulance MBANSs and MSS-CGC

| Victim characteristics | Units | CGC MES | MBANSs |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1.23 | 3 |
| Receiver noise figure | dB | 4 | 10 |
| Receiver antenna height | m | 1.5 | 1.5 |
| Receiver antenna gain | dBi | 0 | 0 |
| Operating frequency | MHz | 2490 | 2490 |
| N, receiver thermal noise | dBm | -108.93 | -99.1 |
| I/N objective | dB | -12 | 0 |
| Interferer’s characteristics |  | MBANSs | CGC BS Tx |
| e.i.r.p | dBm | 0 | 17 |
| Bandwidth | MHz | 3 | 1.23 |
| BW correction factor | dB | -3.87 | 0 |
| NFD (adjacent band interf) | dB | 0 | 0 |
| Wall attenuation | dB | 7 | 7 |
| Antenna height | m | 1.5 | 30 |
| Minimum path loss | dB | 110.1 | 109.1 |
| Interference distance FSL model | km | 3.05 | 2.72 |
| Interference distance P.452 model with 50% probability | km | 0.91 | 2.71 |

#### Interference from CGC base station to MBANSs

Figure 61: Interference scenario – CGC BS into ambulance MBAN

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB   penetration loss

Wanted signal paths

Interference paths

Victim MES receiver

CGC BS

Table 81: Interference from CGC BS to ambulance MBANS - Settings and results

| Simulation input/output parameters | Settings / Results |
| --- | --- |
| Victim Link (VLK): MBANS |  |
| VLK frequency | 2490 MHz |
| VLK bandwidth | 3 MHz |
| VLT → VLR path | FSL (user-defined radius, 3 m) |
| Interfering Link (ILK): CGC link (BS to MES) |  |
| ILK frequency | 2490 MHz |
| ILK bandwidth | 5 MHz |
| ILT Tx power | 43 dBm |
| ILT peak antenna gain | 19 dBi |
| ILT density | 0.02/km2 |
| ILT probability of transmission | 1 |
| ILT → VLR interfering path | Extended Hata, urban, outdoor→indoor, above roof |
| ILT → VLR positioning mode | Uniform density |
| ILT → VLR number of active transmitters | 1 |
| Simulation results |  |
| dRSS, dBm (Std.dev., dB) | -46.8 (4.44) |
| iRSSunwanted, dBm (Std.dev., dB) | -97.84 (12.82) |
| Probability of interference (%) (C/I = 15 dB) | 0.9 |

#### Interference from MBANSs to Globalstar MES receivers in CGC mode

This interference situation is believed to be equivalent to that analyzed in 6.3.3.1. The same results apply, namely 1.4% probability of interference.

### Summary CGC

[

Healthcare facility MBANSs:

According to the results presented above, in worst-case situations—modelled with MCL calculations—the interference distance between CGC and MBANSs is 1.5 km - 6.8 km (MBANS as interferer) and 1.4 km (CGC BS as interferer). The simulation-based average-case analysis yielded 1.8% - 3.5% interference probability from MBANSs into CGC (i.e. MSS terminal in CGC mode). In the other coexistence direction, the probability of interference from CGC into MBANSs is 0.4%.

CGC systems operating in the 2483.5 — 2500 MHz range and healthcare facility MBANSs operating in the same range appear to be compatible.

Home MBANSs:

In worst-case situations—modelled with MCL calculations—the interference distance between CGC and MBANSs is 4.4 km - 31 km (MBANS as interferer) and 5.2 km – 6.1 km (CGC BS as interferer). The simulation-based average-case analysis yielded <1% interference probability from MBANSs into CGC (i.e. MSS terminal in CGC mode). In the other coexistence direction, the probability of interference from CGC into MBANSs is 0.3%.

CGC systems operating in the 2483.5 — 2500 MHz range and home MBANSs operating in the same range appear to be compatible.

Ambulance MBANSs:

In worst-case situations—modelled with MCL calculations—the interference distance between CGC and MBANSs is 900 m – 3.1 km (MBANS as interferer) and 2.7 km (CGC BS as interferer). The simulation-based average-case analysis yielded <1% interference probability from MBANSs into CGC (i.e. MSS terminal in CGC mode). In the other coexistence direction, the probability of interference from CGC into MBANSs is 1.4%.

CGC systems operating in the 2483.5 — 2500 MHz range and ambulance MBANSs operating in the same range appear to be compatible.

]

## RDSS/RNSS (Galileo)

The upgrading of Radiodetermination Satellite Service (RDSS) in this frequency band to primary status on a global basis has been agreed by WRC-12. The approved global primary allocation is intended to facilitate new navigation signals for the next generation of Galileo satellites in subject frequency band. The band 2483.5 – 2500 MHz, because of its proximity to the mobile service allocations above 2.5 GHz, may offer attractive synergies of Radionavigation-satellite (RNSS) with terrestrial mobile systems due to improved antenna efficiencies and use of shared hardware not possible with other RNSS bands. Although upgrading of allocation is conditional on the new service being able to prove its compatibility with other primary services already existing in the band, it was regarded as necessary to consider the co-existence between MBANSs and RNSS.

The interference scenario with RDSS/RNSS is very similar to the scenario with MSS discussed in Section 6.3. RNSS would also be operated in downlink mode only, and interference from the satellite signal to MBANS receivers may be disregarded. For that reason, only the interference from an MBANS transmitter to a RNSS receiver is considered. Another similarity with the MSS case is that also RNSS victim receivers are supposed to be operated outdoors only.

The relevant technical RNSS-receiver parameters used for the study were collected from ECC Rep 150[]  ECC Rep 165 [] and ECC Rep149 [] were also consulted.

Table 82: RNSS (Galileo) system parameters

|  | Parameter | Value |  |
| --- | --- | --- | --- |
| Satellite Transmitter | Bandwidth (MHz) | 1 (reference value) (from ECCREP165[] and ECCREP149[])) | 4 (from range of values in ECCREP150[]) |
|  | Max p.f.d. on ground (dBm/MHz) | -116 |  |
| Mobile Receiver | Maximum interference value | -116 (dBm/MHz) | -110 (dBm/4MHz) |
|  | Antenna height (m) | 1.5 |  |
|  | Antenna gain (dBi) | 0 |  |

### Healthcare facility MBANSs

Table 84: MCL calculation in 2483,5 – 2500 MHz band between healthcare MBANSs and RNSS

| Victim characteristics | Units | RNSS Receiver |  |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1 | 4 |
| Receiver antenna height | m | 1.5 |  |
| Receiver antenna gain | dBi | 0 |  |
| Operating frequency | MHz | 2490 |  |
| Maximum interference value | dBm/BWv | -116 | -110 |
| Interferer’s characteristics |  | MBANSs |  |
| e.i.r.p | dBm | 0 |  |
| Bandwidth | MHz | 3 |  |
| BW correction factor | dB | -4.77 | 0 |
| NFD (adjacent band interf) | dB | 0 |  |
| Wall attenuation | dB | 13 |  |
| Antenna height | m | 1.5 |  |
| Minimum path loss | dB | 98.2 | 97 |
| Interference distance FSL model (propagation exponent n=2) | km | 0.78 | 0.68 |
| Interference distance FSL model (propagation exponent n=3) | km | 0.08 | 0.08 |

#### Interference from MBANSs to RNSS receivers (Galileo)

The considered interference scenario is depicted in Figure 59 and the simulation settings and results are summarized in Table 89.

MBANS

sensor

MBANS

hub

e.i.r.p= 0 dBm

13 dB building penetration loss

Healthcare facility

Wanted signal paths

Interference paths

Victim RNSS receiver

Figure 62: Interference scenario – healthcare facility MBANS into RNSS receiver

Table 85: Interference from healthcare facility MBANS to RNSS - Settings and results

| Simulation input/output parameters | Settings / Results |  |  |  |
| --- | --- | --- | --- | --- |
| Victim Link (VLK): RNSS DL (satellite to mobile receiver) |  |  |  |  |
| VLK frequency | 2490 MHz |  |  |  |
| VLK bandwidth | 1 MHz |  | 4 MHz |  |
| VLR Noise floor | -116 dBm/MHz |  | -110 dBm/4MHz |  |
| Interfering Link (ILK): MBANS |  |  |  |  |
| ILK frequency | 2490 MHz |  |  |  |
| ILK bandwidth | 3 MHz |  |  |  |
| ILT Tx power | 0 dBm |  |  |  |
| ILT density | 40/km2 |  |  |  |
| ILT probability of transmission | 0.1 |  |  |  |
| ILT → VLR interfering path | Extended Hata, indoor→outdoor, below roof |  |  |  |
|  | Urban | Suburban | Urban | Suburban |
| ILT → VLR positioning mode | Closest interferer |  |  |  |
| Simulation results |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |  |  |  |
| iRSSunwanted, dBm (Std.dev., dB) | -164.04 (19.86) | -152.24 (18.88) | -159.38 (19.94) | -147.46 (18.88) |
| Probability of interference (%) (I/N = 0 dB) | 2.6 | 4.5 | 2.6 | 4.2 |

### Home MBANSs

Table 86: MCL calculation in 2483.5 – 2500 MHz band between home MBANSs and RNSS

| Victim characteristics | Units | RNSS Receiver |  |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1 | 4 |
| Receiver antenna height | m | 1.5 |  |
| Receiver antenna gain | dBi | 0 |  |
| Operating frequency | MHz | 2490 |  |
| Maximum interference value | dBm/BWv | -116 | -110 |
| Interferer’s characteristics |  | MBANSs |  |
| e.i.r.p | dBm | 13 |  |
| Bandwidth | MHz | 3 |  |
| BW correction factor | dB | -4.77 | 0 |
| NFD (adjacent band interf) | dB | 0 |  |
| Wall attenuation (Note 1) | dB | 0 (outdoor) |  |
| Antenna height | m | 1.5 |  |
| Minimum path loss | dB | 124.2 | 123 |
| Interference distance FSL model (propagation exponent n=2) | km | 15.6 | 13.51 |
| Interference distance FSL model (propagation exponent n=3) | km | 0.62 | 0.57 |

Note 1: Worst case configuration assumed, in which MBANS transmitter or receiver is outdoors.

Interference from MBANSs to RNSS receivers (Galileo)

MBANS

sensor

MBANS

hub

e.i.r.p= 13 dBm

13 dB building penetration loss

Home

Wanted signal paths

Interference paths

Victim RNSS receiver

Figure 63: Interference scenario – home MBANS into RNSS receiver

Table 87: Interference from home MBANS to RNSS - Settings and results

| Simulation input/output parameters | Settings / Results |  |  |  |
| --- | --- | --- | --- | --- |
| Victim Link (VLK): RNSS DL (satellite to mobile receiver) |  |  |  |  |
| VLK frequency | 2490 MHz |  |  |  |
| VLK bandwidth | 1 MHz |  | 4 MHz |  |
| VLR Noise floor | -116 dBm/MHz |  | -110 dBm/4MHz |  |
| Interfering Link (ILK): Home MBANS |  |  |  |  |
| ILK frequency | 2490 MHz |  |  |  |
| ILK bandwidth | 3 MHz |  |  |  |
| ILT Tx power | 13 dBm |  |  |  |
| ILT density | 10/km2 |  |  |  |
| ILT probability of transmission | 0.02 |  |  |  |
| ILT → VLR interfering path | Extended Hata, indoor→outdoor, below roof |  |  |  |
|  | Urban | Suburban | Urban | Suburban |
| ILT → VLR positioning mode | Uniform |  |  |  |
| Simulation results |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |  |  |  |
| iRSSunwanted, dBm (Std.dev., dB) | -166.56 (14.13) | -153.25 (13.8) | -160.41 (14.15) | -148.27 (13.93) |
| Probability of interference (%) (I/N = 0 dB) | 0.7 | 1.5 | 0.6 | 1.5 |

### Ambulance MBANSs

Table 88: MCL calculation in 2483.5 – 2500 MHz band between ambulance MBANSs and RNSS

| Victim characteristics | Units | RNSS Receiver |  |
| --- | --- | --- | --- |
| Receiver bandwidth | MHz | 1 | 4 |
| Receiver antenna height | m | 1.5 |  |
| Receiver antenna gain | dBi | 0 |  |
| Operating frequency | MHz | 2490 |  |
| Maximum interference value | dBm/BWv | -116 | -110 |
| Interferer’s characteristics |  | Ambulance MBANSs |  |
| e.i.r.p | dBm | 0 |  |
| Bandwidth | MHz | 3 |  |
| BW correction factor | dB | -4.77 | 0 |
| NFD (adjacent band interf) | dB | 0 |  |
| Wall attenuation | dB | 7 |  |
| Antenna height | m | 1.5 |  |
| Minimum path loss | dB | 104.2 | 103 |
| Interference distance FSL model (propagation exponent n=2) | km | 1.56 | 1.35 |
| Interference distance FSL model (propagation exponent n=3) | km | 0.13 | 0.12 |

#### Interference from MBANSs to RNSS receivers (Galileo)

The considered interference scenario is depicted in Figure 61 and the simulation settings and results are summarized in Table 93.

MBANS

hub

MBANS

sensor

e.i.r.p= 0 dBm

7 dB   penetration loss

Wanted signal paths

Interference paths

Victim RNSS receiver

Figure 64: Interference scenario ambulance MBANS into RNSS receiver

Table 89: Interference from MBANS (ambulance mode) to RNSS - Settings and results

| Simulation input/output parameters | Settings / Results |  |  |  |
| --- | --- | --- | --- | --- |
| Victim Link (VLK): RNSS DL (satellite to mobile receiver) |  |  |  |  |
| VLK frequency | 2490 MHz |  |  |  |
| VLK bandwidth | 1 MHz |  | 4 MHz |  |
| VLR Noise floor | -116 dBm/MHz |  | -110 dBm/4MHz |  |
| Interfering Link (ILK): ambulance MBANS |  |  |  |  |
| ILK frequency | 2490 MHz |  |  |  |
| ILK bandwidth | 3 MHz |  |  |  |
| ILT Tx power | 0 dBm |  |  |  |
| ILT density | 5/km2 |  |  |  |
| ILT probability of transmission | 0.1 |  |  |  |
| ILT → VLR interfering path | Extended Hata, indoor→outdoor, below roof |  |  |  |
|  | Urban | Suburban | Urban | Suburban |
| ILT → VLR positioning mode | Uniform |  |  |  |
| Simulation results |  |  |  |  |
| dRSS, dBm (Std.dev., dB) | -100 (0) |  |  |  |
| iRSSunwanted, dBm (Std.dev., dB) | -164.2 (13.88) | -150 (13.08) | -157.51 (13.71) | -145.25 (13.36) |
| Probability of interference (%) (I/N = 0 dB) | 1.1 | 2.2 | 1.0 | 2.2 |

### Summary RDSS/RNSS

[

According to the results presented above, in worst-case situations—modelled with MCL calculations—the interference distance between RNSS and MBANSs is 80 m - 780 m (healthcare facility MBANS), 570 m - 15 km (home MBANS), and 120 m – 1.6 km (ambulance MBANS). The simulation-based average-case analysis yielded probability of interference levels of ≤4.5% from healthcare facility MBANSs, ≤1.5% from home MBANSs, and ≤2.2% from ambulance MBANSs.

RDSS/RNSS systems operating in the 2483.5 — 2500 MHz range and MBANSs operating in the same range appear to be compatible.

]

## SAP/SAB (ENG/OB) video links

The SAP/SAB (ENG/OB) tuning range extends up to 2500 MHz. Although ECC Report 002 [ ] shows that most of the SAP/SAB use is concentrated below 2400 MHz, some CEPT countries have reported using the entire tuning range for SAP/SAB operations, which leads to the consideration of SAP/SAB as incumbent in the 2483.5 – 2500 MHz band. The interference situation is very similar to that studied in Section 4.4 for the 2300 – 2400 MHz band. The only difference is the operation frequency, which—after verification with test simulations—does not have a significant influence on the interference probability results. Therefore the same conclusions as in Section 4.4 apply to SAP/SAB (ENG/OB) in the 2483.5 – 2500 MHz band.

## ADJACENT Wideband data transmission systems

This section studies the compatibility between MBANSs operating in the 2483.5 — 2500 MHz range and wideband data transmission systems–represented by WLAN– operating in the adjacent 2400 — 2483.5 MHz range.  Due to the equivalence to the co-channel compatibility study in Section 5.1, this section is presented in compact form. More details and considerations can be found in Section 5.1.

Based on the WLAN standard [] the last channel of wideband data transmission on the upper side of the band ends at 2482 MHz. That results in 1.5 MHz buffer between WLAN and the possible MBANs implementation in the band studied in this section (2483.5 – 2500 MHz). However, it was considered appropriate to also study possible interference between these technologies in these two adjacent bands. The MCL calculations for 20 MHz WLAN channel bandwidth are presented in Table 91. The same approach as in the co-channel compatibility analysis in Section 5.1 was used.

Table 91: Adjacent MCL calculation in 2483.5 MHz band limits between MBANSs and WLAN

| Victim characteristics | Units | WLAN 802.11 AP | WLAN 802.11 AP | MBANSs |
| --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 20 | 20 | 3 |
| Receiver noise figure | dB | 10 | 10 | 10 |
| Receiver antenna height | m | 3 | 3 | 1.5 |
| Receiver antenna gain | dBi | 5 | 5 | 0 |
| Operating frequency | MHz | 2472 | 2472 | 2485 |
| N, receiver thermal noise | dBm | -90.82 | -90.82 | -99.1 |
| I/N objective | dB | 0 | 0 | 0 |
| Interferer’s characteristics |  | MBANSs (healthcare facility) | MBANSs (home) | WLAN 802.11 AP |
| e.i.r.p | dBm | 0 | 13 | 17 |
| Bandwidth | MHz | 3 | 3 | 20 |
| BW correction factor | dB | 0 | 0 | -8.24 |
| NFD (adjacent band interf.) | dB | -25 | -25 | -40 |
| Antenna height | m | 1.5 | 1.5 | 3 |
| Minimum path loss | dB | 70.8 | 83.8 | 67.8 |
| Interference distance FSL model | km | 0.03 | 0.15 | 0.02 |

### Healthcare facility MBANSs

The SEAMCAT simulations are summarized in the following table. The study approach is identical to that Section 5.1.1. The parameter and scenario choices are described in that section of the report. For obtaining the interference probability in SEAMCAT, the combination of unwanted and blocking response was used. A value of 30 dB blocking response was used for MBANSs.

Table 92: Interference in 2483.5 MHz band limits between healthcare facility MBANSs and WLAN – Simulation results

| Interferer | Victim | Interferer central frequency (MHz) | Victim central frequency (MHz) | Simulation parameters | Interference probability / bitrate degradation |
| --- | --- | --- | --- | --- | --- |
| MBANS | WLAN (20 MHz) | 2485 | 2472 | See Section 5.1.1.1 | 0.1 % |
|  | WLAN (40 MHz) |  | 2462 |  | 0 % |
| WLAN (20 MHz) | MBANS | 2472 | 2485 | See Section 5.1.1.2 | 4.2 % |
| WLAN (40 MHz) |  | 2462 |  |  | 3.9 % |

### Home MBANSs

The SEAMCAT simulations are summarized in the following table. The study approach is identical to that Section 5.1.1. The parameter and scenario choices are described in that section of the report. For obtaining the interference probability in SEAMCAT, the combination of unwanted and blocking response was used. A value of 30 dB blocking response was used for MBANSs.

Table 93: Interference in 2483.5 MHz band limits between home MBANSs and WLAN – Simulation results

| Interferer | Victim | Interferer central frequency (MHz) | Victim central frequency (MHz) | Simulation parameters | Interference probability / bitrate degradation |
| --- | --- | --- | --- | --- | --- |
| MBANS | WLAN (20 MHz) | 2485 | 2472 | See Section 5.1.2.1 | 0 % |
|  | WLAN (40 MHz) |  | 2462 |  | 0.1 % |
| WLAN (20 MHz) | MBANS | 2472 | 2485 | See Section 5.1.2.2 | 3.1 % |
| WLAN (40 MHz) |  | 2462 |  |  | 3 % |

### Summary ADJACENT wideband data transmission systems

[

All simulation-based average-case analyses yielded ≤0.1% probability of interference from MBANSs. In the other coexistence direction, the obtained results range from 3% to 4.2% probability of interference from wideband data transmission systems.

Wideband data transmission systems operating in the 2400 — 2483.5 MHz range and MBANSs in healthcare facilities, homes, and ambulances operating in the 2483.5 — 2500 MHz range are compatible.

]

## Adjacent broadband wireless systems (bws)

[ENTIRE SECTION 6.8 NEEDS REVISION AFTER THE REVISION OF SECTION 4.3 IS FINALIZED]

The band 2500-2690 MHz is another band where LTE can be utilized. Since the involved power levels are quite high and there are currently no guard bands specified for the LTE, it is important to study also possible adjacent band interferences. Therefore the possible adjacent band interferences between this LTE band and the considered 2483.5-2500 MHz band is analysed.

All of the parameters of both systems are the same as in Section 4.3, with the only difference being the frequency offset. This leads also to a corresponding reduction of the power levels due to the shape of the emission masks. More specifically we studied the worst case, where both technologies are utilizing the edge of their allocated (or presumably designated) spectrum without any guard bands. The ACLR values of LTE devices were taken from the LTE standard.  The ACLR value for MBANS is taken as the average of the adjacent band emissions in the considered spectrum of 10 MHz for LTE.

### Healthcare facility MBANSs

The MCL calculations are presented in Table 24. The simulation results are mentioned in Table 25.

Table 24: Adjacent MCL calculation in 2500 MHz band limits between healthcare facility MBANSs and LTE

| Victim characteristics | Units | LTE DL | LTE UL |  | MBANSs |  |
| --- | --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 10 | 10 |  | 3 |  |
| Receiver noise figure | dB | 9 | 5 |  | 10 |  |
| Receiver antenna height | m | 1,5 | 30 |  | 1,5 |  |
| Receiver antenna gain | dBi | 0 | 14 |  | 0 |  |
| Operating frequency | MHz | 2505 | 2505 |  | 2498,5 |  |
| N, receiver thermal noise | dBm | -94,8 | -98,8 |  | -99,1 |  |
| I/N objective | dB | -6 | -6 |  | 0 |  |
| Interferer’s characteristics |  | MBANSs |  |  | LTE DL | LTE UL |
| e.i.r.p | dBm | 0 |  |  | 60 | 23 |
| Bandwidth | MHz | 3 |  |  | 10 |  |
| BW correction factor | dB | 0 |  |  | -5,2 |  |
| NFD (adjacent band interf) | dB | -25 |  |  | -45 | -30 |
| Wall attenuation | dB | 0 (LTE UE indoor) /13 (LTE UE outdoor) |  | 13 | 13 | 0 |
| Antenna height | m | 1,5 |  |  | 30 | 1,5 |
| Minimum path loss | dB | 75,8 / 62,8 |  | 80,8 | 95,8 | 86,8 |
| Interference distance FSL model | km | 0,06 / 0,01 |  | 0,10 | 0,59 | 0,21 |

Table 25: Interference in 2500 MHz band limits between healthcare facility MBANSs and LTE – Simulation results

| Interferer | Victim | Interferer central frequency (MHz) | Victim central frequency (MHz) | Simulation parameters | Interference probability / bitrate degradation (urban/suburban) |
| --- | --- | --- | --- | --- | --- |
| LTE DL | MBANS | 2505 | 2498.5 | See Section 4.3.1.1 | 0.1 % / 0.2% |
| MBANS | LTE DL | 2498.5 | 2505 | See Section 4.3.1.2 | 0.1 % |
| LTE UL | MBANS | 2505 | 2498.5 | See Section 4.3.1.3 | 1.4 % / 0% |
| MBANS | LTE UL | 2498.5 | 2505 | See Section 4.3.1.4 | 0.8 % /1.9% |

### Home MBANSs

The MCL calculations are presented in Table 26. The simulation results are mentioned in Table 27.

Table 26: Adjacent MCL calculation in 2500 MHz band limits between home MBANSs and LTE

| Victim characteristics | Units | LTE DL | LTE UL |  | MBANSs |  |
| --- | --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 10 | 10 |  | 3 |  |
| Receiver noise figure | dB | 9 | 5 |  | 10 |  |
| Receiver antenna height | m | 1,5 | 30 |  | 1,5 |  |
| Receiver antenna gain | dBi | 0 | 14 |  | 0 |  |
| Operating frequency | MHz | 2505 | 2505 |  | 2498,5 |  |
| N, receiver thermal noise | dBm | -94,8 | -98,8 |  | -99,1 |  |
| I/N objective | dB | -6 | -6 |  | 0 |  |
| Interferer’s characteristics |  | MBANSs |  |  | LTE DL | LTE UL |
| e.i.r.p | dBm | 13 |  |  | 60 | (Note 1) |
| Bandwidth | MHz | 3 |  |  | 10 |  |
| BW correction factor | dB | 0 |  |  | -5,2 |  |
| NFD (adjacent band interf) | dB | -25 |  |  | -45 |  |
| Wall attenuation | dB | 0 (both indoor or outdoor) /13 (complementary scenarios) |  |  | 0 |  |
| Antenna height | m | 1,5 |  |  | 30 |  |
| Minimum path loss | dB | 88,8 / 75,8 |  | 106,8 / 93,8 | 108,8 |  |
| Interference distance FSL model | km | 0,26 / 0,06 |  | 2,09 / 0,47 | 2,64 |  |

Note 1: same as for healthcare facility MBANSs. See Table 24.

Table 27: Interference in 2400 MHz band limits between home MBANSs and LTE – Simulation results

| Interferer | Victim | Interferer central frequency (MHz) | Victim central frequency (MHz) | Simulation parameters | Interference probability / bitrate degradation |
| --- | --- | --- | --- | --- | --- |
| LTE DL | MBANS | 2505 | 2498.5 | See Section 4.3.2.1 | 0.1 % |
| MBANS | LTE DL | 2498.5 | 2505 | See Section 4.3.2.2 | 0 % (all cases) / 0.1 % (worst case) |
| LTE UL | MBANS | 2505 | 2498.5 | See Section 4.3.2.3 | 1.1 %/ 0 % |
| MBANS | LTE UL | 2498.5 | 2505 | See Section 4.3.2.4 | 0.7 - 1.7 % / 1.8-10.1 % (MBANS indoor-outdoor) |

### Ambulance MBANSs

The MCL calculations are presented in Table 28. The simulation results are presented in Table 29.

Table 28: Adjacent MCL calculation in 2400 MHz band limits between home MBANSs and LTE

| Victim characteristics | Units | LTE DL | LTE UL |  | MBANSs |  |
| --- | --- | --- | --- | --- | --- | --- |
| Receiver bandwidth | MHz | 10 | 10 |  | 3 |  |
| Receiver noise figure | dB | 9 | 5 |  | 10 |  |
| Receiver antenna height | m | 1,5 | 30 |  | 1,5 |  |
| Receiver antenna gain | dBi | 0 | 14 |  | 0 |  |
| Operating frequency | MHz | 2505 | 2505 |  | 2498,5 |  |
| N, receiver thermal noise | dBm | -94,8 | -98,8 |  | -99,1 |  |
| I/N objective | dB | -6 | -6 |  | 0 |  |
| Interferer’s characteristics |  | MBANSs |  |  | LTE DL | LTE UL |
| e.i.r.p | dBm | 0 |  |  | 60 | 23 |
| Bandwidth | MHz | 3 |  |  | 10 |  |
| BW correction factor | dB | 0 |  |  | -5,2 |  |
| NFD (adjacent band interf) | dB | -25 |  |  | -45 | -30 |
| Wall attenuation | dB | 7 |  |  | 7 |  |
| Antenna height | m | 1,5 |  |  | 30 | 1,5 |
| Minimum path loss | dB | 68,8 |  | 86,8 | 101,8 | 79,8 |
| Interference distance FSL model | km | 0,03 |  | 0,21 | 1,18 | 0,09 |

Table 29: Interference in 2400 MHz band limits between home MBANSs and LTE – Simulation results

| Interferer | Victim | Interferer central frequency (MHz) | Victim central frequency (MHz) | Simulation parameters | Interference probability / bitrate degradation |
| --- | --- | --- | --- | --- | --- |
| LTE DL | MBANS | 2505 | 2498.5 | See Section 4.3.3.1 | 0.1 % |
| MBANS | LTE DL | 2498.5 | 2505 | See Section 4.3.3.2 | 0 % |
| LTE UL | MBANS | 2505 | 2498.5 | See Section 4.3.3.3 | 3.7 % /0.6 % |
| MBANS | LTE UL | 2498.5 | 2505 | See Section 4.3.3.4 | 0.1 % / 0.2 % |

### Summary adjacent broadband wireless systems

[no issues?]

# Conclusions

[Comment MATTIAS:

The duty cycle was assumed with 10 % for all compatibility scenarios. This important information should be given here in the conclusion. …

Compatibility study with railway applications are skipped because of limitation of frequency range (2400 – 2440 MHz). This information should be included in the coclusion…..

Compatibility study with RFID applications are skipped because of limitation of frequency range (2400 – 2440 MHz). This information should be included in the coclusion.]

[mention that healthcare facility MBANS use a means to prevent outdoor usage]

[Ralf: Some informations from TR 101 557 are possibly worth to be noted here:

- Co-existence possibilities: Interference issues between MBANSs and aeronautical telemetry systems (ATS) in the band 2360-2400 MHz can be addressed by avoiding MBANS operation in the proximity of ATS installations, since there are very few of these that use the 2360-2400 MHz spectrum. These proposals are already well developed in the US FCC discussions, and it is recognized that MBANSs can effectively meet the ITU M.1459 recommendations for interference protection [i.16]. Adaptive frequency selection, listen-before-talk, adaptive power control and other features may possibly enable opportunistic low-power, short-range MBANSs to achieve harmonized coexistence with ATS and amateur radios. It is possible that countries in Europe may issue spectrum authorisations for mobile broadband applications including IMT in the 2300-2400 MHz band based on unpaired operation in contiguously aggregated 5 MHz blocks.

- APC may be used by MBANS transmitters

Some questions to those points

]

[ PRELIMINARY CONCLUSION OVERVIEW

|  |  | Healthcare facility MBANSs | Ambulance MBANSs | Home MBANSs |
| --- | --- | --- | --- | --- |
| Tx-power, DC |  | 1 mW, 10% DC | 1 mW, 10% DC | 20 mW, 2% DC |
| Description of MBANS usage type and restriction |  | Only indoor, within healthcare facility | Only inside vehicle | Primarily indoor, within patient home. Occasionally outdoor, few meters from home |
| 2300 – 2400 MHz | ATS | compatible, with 0.7 km geographical separation to ATS Earth station | not compatible | not compatible |
|  | Amateur | compatible | compatible | not compatible |
|  | BWS | geo-separation to BWS of x km ? | ? | ? |
|  | SAP/SAB | compatible | not compatible | not compatible |
|  | Wideband Data (adjacent) | compatible | compatible | compatible |
|  | Deep Space SRS (adjacent) | compatible, with 1.5 km geographical separation to SRS Earth station | not compatible | not compatible |
| 2400 – 2483.5 MHz | Wideband Data | not compatible | compatible | not compatible |
|  | ISM | – | – | – |
|  | Non-specific SRDs | – | – | – |
|  | Railway | Compatible, if MBANSs operate in 2400 – 2440 MHz range |  |  |
|  | RFID |  |  |  |
|  | Amateur and Amateur Satellite | compatible | compatible | not compatible |
|  | BWS (adjacent) | ? | ? | ? |
| 2483.5 – 2500 MHz | LP-AMI | not compatible | compatible | compatible |
|  | ISM | – | – | – |
|  | MSS | compatible | compatible | compatible |
|  | CGC | compatible | compatible | compatible |
|  | RNSS/RDSS | compatible | compatible | compatible |
|  | SAP/SAB | compatible | not compatible | not compatible |
|  | Wideband Data (adjacent) | compatible | compatible | compatible |
|  | BWS (adjacent) | ? | ? | ? |

]

heading (style: ECC annex - heading1)

Body text (style: ECC Paragraph)

heading 2 (style: ECC annex heading2)

List of referenceS

- Reference one (style: reference)

Reference two

Etc.

Any bandwidth in the indicated range is applicable. Interference distance result is not affected by choice.